Owner’s Manual for the

PulseBlasterDDS™ - ISA Board

Complete Digital Excitation System:

Direct Digital Synthesis (DDS),

Arbitrary Waveforms, and TTL

SpinCore Technologies, Inc.

3525 NW 67th Avenue

Gainesville, Florida 32653, USA

Phone: (352)-271-7383

http://www.spincore.com

Congratulations and THANK YOU for choosing a design from SpinCore Technologies, Inc.  We appreciate your business.  At SpinCore we try to fully support the needs of our customers, so if you ever need assistance please contact us and we will strive to provide the necessary help.

© 2000-2002 SpinCore Technologies, Inc. All rights reserved. SpinCore Technologies, Inc. reserves the right to make changes to the product(s) or information herein without notice. PulseBlasterDDS™, PulseBlaster™, SpinCore, and the SpinCore Technologies, Inc. logo are trademarks of SpinCore Technologies, Inc. All other trademarks are the property of their respective owners.

SpinCore Technologies, Inc. makes every effort to verify the correct operation of the equipment.  This equipment should NOT, however, be used in system where the failure of a SpinCore device will cause serious damage to other equipment or harm to a person.

Contents

| Section I: Introduction # 1 Product Overview 2 Board Architecture Section II: PulseBlasterDDS Design 1 Programming Architecture 2 Quick Installation Guide 3 Header/Jumper Information 4 ISA Readback Feature Section III: Programming Notes: Using C Functions (Provided) - Provided C Functions: Description - Example Use of C Functions Appendix I: Programming Notes: Sample C Program Appendix II Programming Notes: Writing Directly to I/O Ports | 4 5 8 12 13 17 19 22 23 26 |
| --- | --- |

Section I: Introduction

1. Product Overview

The PulseBlasterDDS™ series of Intelligent Pattern and Waveform Generation boards from SpinCore Technologies, Inc., couples SpinCore’s unique Intelligent Pattern Generation and Timing Processor Core, dubbed PulseBlaster™, with Direct Digital Synthesis (DDS) AND Arbitrary Waveform Generation (AWG) technology for use in system control and waveform generation.

The PulseBlaster’s processor core provides all the necessary timing control signals required for overall system control and waveform synchronization.   By adding DDS and arbitrary waveform generation features, PulseBlasterDDS™ can now provide not only TTL but also digitally synthesized analog output signals, meeting high-performance and high-precision complex excitation/stimuli needs of demanding users.

PulseBlasterDDS™ provides users the ability to control their systems through the generation of fully synchronized (digital and analog) excitation waveforms from a small form factor PC board, providing users a compelling price/performance proposition unmatched by any other device on the market today.  Figure 1 presents sample capabilities of the board.

Figure 1. Sample PulseBlasterDDS(tm) output capabilities.

2. Board Architecture

Block diagram

Figure 2 presents the general architecture of the PulseBlasterDDS™ system.  The two major building blocks are the DDS Core and the Pulse Programming and Timing Processor Core (PP Core). The DDS Core contains a numerically controlled oscillator and has 16 programmable frequency registers that are under the pulse program control.  Prior to gating, the DDS signal can be phase offset by one of 16 programmable phase registers for each of the two independent output channels.  An optional modulator can be placed in the path of one of the output channels.  The PP Core controls the timing of the gating pulses and provides the necessary control signals for frequency and phase registers.  The DDS and PP cores have been integrated onto a single silicon chip.  High performance DAC chips and high current output amplifiers complement the design.

Fig. 2. PulseBlasterDDS(tm) board architecture.  The DDS and PP cores have been integrated onto a single silicon chip.

Output signals

The PulseBlasterDDS™ product comes with up to two analog output channels that can be configured to output radio-frequency (RF/IF) waveforms, arbitrary waveforms, or a combination of both, and up to 10 digital output signal lines.  The frequency and phase of the RF signals generated by the DDS output channels are under the complete control of the user and are specified through software programming.  The arbitrary waveform patterns are also under the control of the user and are stored in on-board memory.  PulseBlasterDDS™ also provides the ability to gate the output of the DDS channels allowing for independent pulsed RF operation of the output channels.  With digital sampling rate of 100 MHz, analog signals up to approx. 50 MHz can be generated.  Both analog output signals are available on on-board SMA connectors.  The output impedance of the analog signals is 50 ohms.

The individually controlled digital (TTL/CMOS) output bits are capable of delivering +-25 mA per bit.  Output lines, available on the PC bracket-mounted DB-25 connector can be set to either the 5 V or 3.3V I/O TTL logic standard.

##### Timing characteristics

PulseBlasterDDS™’ timing controller can accept either an internal (on-board) crystal oscillator or an external frequency source of up to 100 MHz. The innovative architecture of the timing controller allows the processing of either simple timing instructions (delays of up to 2^32 = 4,294,967,296 clock cycles), or double-length timing instructions (up to 2^52 clock cycles long - nearly 2 years with a 100 MHz clock!).  Regardless of the type of timing instruction, the timing resolution remains constant for any delay – just one clock period (e.g., 10 ns for a 100 MHz clock).

The core timing controller has a very short minimum delay cycle – only five clock periods for internal memory (512 words) models.  This translates to a 50 ns pulse/delay/interval with a 100 MHz clock.  The external memory models (up to 32k words) have a nine-clock-period minimum instruction cycle.

##### Instruction set

PulseBlasterDDS™’ design features a set of commands for highly flexible program flow control. The micro-programmed controller allows for programs to include branches, subroutines, and loops at up to 8 nested levels – all this to assist the user in creating dense pulse programs that cycle through repetitious events, especially useful in numerous multidimensional spectroscopy and imaging applications.

##### External triggering

PulseBlasterDDS™ can be triggered and/or reset externally via dedicated hardware lines.  The two separate lines combine the convenience of triggering (e.g., in cardiac gating) with the safety of the "stop/reset" line.  The required control signals are “active low” (or short to ground).

##### Summary

PulseBlasterDDS™ is a versatile, high-performance pulse/pattern TTL and RF/IF/AWG generator operating at speeds of up to 100 MHz and capable of generating pulses/delays/intervals ranging from 50 ns to over 2 years per instruction. It can accommodate pulse programs with highly flexible control commands of up to 32k program words. Its high-current output logic bits are independently controlled and the output bits are 5/3.3 V user-selectable.  Up to two analog channels (50 ohm output impedance) are available from a single board.

##### Summary Specifications

DDS Section (Model DDS-100):

- 100 MHz reference clock oscillator (other frequencies upon request)

- 0.047 Hz frequency resolution (32 bits)

- 16 loadable frequency registers for agile frequency modulation/switching/selection (32 bits each)

- 16 loadable phase-offset registers for agile phase modulation/switching/selection (12 bits each)

- 0.09 degree phase resolution (12 bits)

- 40 ns phase switching latency

- 40 ns frequency switching latency (phase continuous)

- 5 MHz max. modulating frequency

- 12 dBm rf output power

- 50 ohm output impedance

- SMA connectors

TTL Section:

- 10 individually controlled digital output lines (TTL levels).

- variable pulses/delays for every TTL line.

- 25 mA output current per TTL line.

- output lines can be combined to increase the max. output current.

Common Parameters (DDS and TTL Sections):

- 90 ns shortest pulse/interval (50 ns for non-rf models at 100 MHz clock frequency).

- 2 years longest pulse/interval.

- 10 ns pulse/interval resolution.

- rf and TTL pulses are synchronized.

- 32k max. memory space.

- external  triggering and reset – TTL levels.

Pulse Program Control Flow (Common):

- loops, nested 8 levels deep.

- 20 bit loop counters (max. 1,048,576 repetitions).

- subroutines, nested 8 levels deep.

- wait for trigger - 80 ns latency.

- 5 MHz max. retriggering frequency.

Section II: PulseBlasterDDS™ Design

1. Programming Architecture

Machine-Word Definition

The PulseBlaster™ pulse timing and control processor implements an 80-bit wide Very-Long Instruction Word (VLIW) architecture. The VLIW memory words have specific bits/fields dedicated to specific purposes, and every word should be viewed as a single instruction of the micro-controller. The maximum number of instructions that can be loaded to on-board memory is 32k. The execution time of instructions can be varied and is under (self) control by one of the fields of the instruction word – the shortest being five clock cycles (for 512 memory-word models) and the longest being 2^52 clock cycles.  All instructions have the same format and bit length, and all bit fields have to be filled.  Figure 3 shows the fields and bit definitions of the 80-bit instruction word.

Bit Definitions for the 80-bit Instruction Word (VLIW)

Output/Control Word   |    Data Field    |     OP Code    |    Delay Count

(24 bits)             (20 bits)          (4 bits)         (32 bits)

Figure 3. Bit definitions of the 80-bit VLIW instruction/memory word.

Breakdown of 80 bit VLIW instruction word

The 80-bit VLIW is broken up into 4 sections

- Output Pattern and Control Word - 24 bits

- Data Field - 20 bits

- OP Code - 4 bits

- Delay Count - 32 bits

Output Pattern and Control Word:

| Instruction Bit # | Function |
| --- | --- |
| 23 | Selects one of 16 Frequency Registers (bit3) |
| 22 | Selects one of 16 Frequency Registers (bit2) |
| 21 | Selects one of 16 Frequency Registers (bit 1) |
| 20 | Selects one of 16 Frequency Registers (bit 0) |
| 19 | Selects one of 16 TX Phase Registers (bit 3) |
| 18 | Selects one of 16 TX Phase Registers (bit 2) |
| 17 | Selects one of 16 TX Phase Registers (bit 1) |
| 16 | Selects one of 16 TX Phase Registers (bit 0) |
| 15 | Selects one of 16 RX Phase Registers (bit 3) |
| 14 | Selects one of 16 RX Phase Registers (bit 2) |
| 13 | Selects one of 16 RX Phase Registers (bit 1) |
| 12 | Selects one of 16 RX Phase Registers (bit 0) |
| 11 | Enables output of TX signal (0 = on, 1 = off) |
| 10 | Enables output of RX signal (0 = on, 1 = off) |
| 9 | Output Connector DB25 pin 19 |
| 8 | Output Connector DB25 pin 7 |
| 7 | Output Connector DB25 pin 8 |
| 6 | Output Connector DB25 pin 21 |
| 5 | Output Connector DB25 pin 22 |
| 4 | Output Connector DB25 pin 10 |
| 3 | Output Connector DB25 pin 11 |
| 2 | Output Connector DB25 pin 24 |
| 1 | Output Connector DB25 pin 25 |
| 0 | Output Connector DB25 pin 13 |

Data Field and Op Code:

The data field's function is dependent on the OpCode, see the table below:

| OpCode # | OpCode Meaning | Data Field used for |
| --- | --- | --- |
| 0 | Continue | Ignored |
| 1 | Stop | Ignored |
| 2 | Loop | Number of desired Loops - 1 |
| 3 | End Loop | Address of Instruction originating loop |
| 4 | Jump to Subroutine | Address of first subroutine instruction |
| 5 | Return From Subroutine | Ignored |
| 6 | Branch | Address of next instruction |
| 7 | Long Delay | Number of desired loops - 2 |
| 8 | Wait | Ignored |

Delay Count:

The value of the Delay Count field determines how long the current instruction should be executed. The length of the delay is the value in this field multiplied by the clock period.  The allowed minimum value of this field is 0x6 for the 32k memory models and 0x2 for the internal-memory models.  The timing controller has a fixed delay of three clock cycles and the value that one enters into the Delay Count field should account for this inherent delay.

WAIT OpCode

The WAIT OpCode, 0x8, can be used to halt the operation of the PulseBlasterDDS until a hardware or software trigger is detected.  To exit from the halt condition, a software or hardware trigger can be used.   Upon the trigger, there will be a latency between the time the trigger was issued and the time the next instruction is executed.   The latency is equal to the delay value entered in the WAIT instruction line plus a fixed delay of 6 clock cycles.   At a 100 MHz clock rate, the maximum retriggering frequency is approx. 5 MHz.

Programming of DDS frequency and phase registers

The registers of the DDS core are programmed directly over the system bus.  The following is an example use of the C functions that were developed at SpinCore to aid in programming the available DDS frequency and phase registers:

// Start programming the DDS frequency registers

start_programming(FREQ_REGS);

// Program the registers in order from 0 to 2 (out of the max. 16 registers)

// Valid units are MHz, kHz, and Hz - default is MHz

set_freq(1*MHz); // Set register 0

set_freq(2*MHz); // Set register 1

set_freq(3*MHz); // Set register 2

// Start programming the DDS phase registers in the TX channel

start_programming(TX_PHASE_REGS);

// Program the registers in order from 0 to 3 (out of the max. 16 registers)

// Units are in degrees

set_phase(0); // Set register 0

set_phase(22.5); // Set register 1

set_phase(45); // Set register 2

set_phase(67.5); // Set register 3

For details on programming the DDS section of the board, please consult Section III of this Manual.

Programming of the 80-bit VLIW instructions with C functions

The 80-bit VLIW machine instructions can easily be programmed with the use of the C functions that were developed at SpinCore.   The general form of the functions for programming the sequence of events is (control/output_pattern, time_of_execution).  The actual instruction format is as follows:

//int pb_inst(int freq, int tx_phase, int tx_output_enable, int rx_phase, int rx_output_enable, int flags, int inst, int inst_data, int length)

Examples:

//Simple instruction: Execute for 10.0, after which Continue to next instruction

pb_inst(0,0,TX_ANALOG_ON,1,RX_ANALOG_ON,0x3FF,CONTINUE,0,10.0*us);

//Branch Instruction: Branch to label "start" in 1 ms

pb_inst(0,0,TX_ANALOG_OFF,1,RX_ANALOG_ON,0x0,BRANCH,start,1*ms);

The first instruction will last 10 us, during which the RF pulses will be generated on both Tx and Rx channels.   The frequency of Tx and Rx RF will be determined by the content of the 0th frequency register, the phase of the Tx pulse will be determined by the content of the 0th phase register of the Tx channel, and the phase of the Rx pulse will be determined by the content of the 1st phase register of the Rx channel.  While the instruction executes, all 10 TTL bits/channels will be at the logical one (the 0x3FF value in the instruction).

The second instruction will last 1 ms, after which the program execution will jump to the instruction labeled “start.”  During the execution of this instruction, the Rx channel will be turned ON and the Tx channel will be turned “OFF,” while the TTL outputs will be at the logical “zero.”

Please consult Section III for more information on programming the PulseBlasterDDS™ with C functions.

2. Quick Installation Guide

PulseBlasterDDS™ boards are ready to use out of the box.  After unpacking, they can be installed on your computer in any available ISA slot.  Please shut down your computer and turn the power off when installing the board, and use a screw to fasten the bracket.

PulseBlasterDDS™ boards are factory pre-configured to operate with the following default settings:

ISA Base Address:		0x340.

Clock Oscillator:	Internal, installed on board; clock frequency as per customer specification

TTL Output levels:		3.3 V

These settings can be changed using on-board jumpers. Please consult the next chapter for details regarding the jumpers’ information and location.

The board can be used on computers running any operating system that supports the Industry Standard Architecture (ISA) bus, including DOS, Windows, QNX, and Linux.  PulseBlasterDDS™ ISA board is NOT plug-and-play and will not be recognized by your operating system.  Programming of the board can be accomplished by directly writing to specific I/O ports.  SpinCore’s and third party drivers are available for protected operating systems (Win2000, NT, XP, etc.).  Section III of this manual describes programming of PulseBlasterDDS™ boards in details.

3. Header/Jumper Information

The PulseBlasterDDS™ ISA board is a configurable system.  It allows the user to set jumpers on several headers on the PC card to select different modes of operation for the device.

Selecting ISA Bus Address: Header J4, Pins 1-2, 3-4, and 5-6.

The Base Addresses that can be specified range from 0x260 to 0x360.  The default, factory pre-set value for the ISA PulseBlaster™ board is 0x340. This value can be changed, via jumpers on Header J4, according to the Table 3.

| Base Address (in Hex) | Jumper Settings – Header J4 Pins 5-6 Pins 3-4 Pins 1-2 |
| --- | --- |
| 300 | ¦ ¦ ¦ |
| 320 | ¦ ¦ : |
| 340 | ¦ : ¦ |
| 260 | ¦ : : |
| 280 | : ¦ ¦ |
| 270 | : ¦ : |
| 290 | : : ¦ |
| 360 | : : : |

Table 1. Board’s ISA-Bus Base Address Selection

(Legend: | jumper across pins, : no jumper)

(Default Value = 0x340, jumpers 1-2, 5-6).

Output Signal Bits  - Connector DB-25 (J10)

The following table lists the output bits for the PulseBlaster™ Pulse / Pattern Generator Board.

Table 2. Output bits and signals of the PulseBlasterDDS™ board.

Bits 9-0 are grouped on the external DB-25 connector (also marked as J10) provided for accessing the signals. The table also lists several additional output signals that are available to the outside world, as described in the next subsection.

Selecting TTL output voltage levels: JPower1 and Jpower2

The output signals are driven by latches/drivers capable of running off a 5.0-V or 3.3-V supply.  The supply voltage for the drivers is selectable.  Table 4, below, lists the configurations for 5.0-V and 3.3-V output driver operation.

Table 3. Output voltage selection

The JPower1 header selects the operating voltage for the output bits 0-15, and JPower2 independently selects the operating voltage for the output bits 16-23.

Using external trigger/reset lines - Header JTrigger.

HW_Trigger is a signal that is pulled high by default.  When a falling edge is detected (e.g., when shorting pins 3-4), it initiates code execution.   This trigger will also restart execution of a program from the beginning of the code if it is detected after the design has reached an idle state.  The idle state could have been created either by reaching the STOP Op Code of a program, or by the detection of the HW_Reset signal.  When the WAIT OpCode is used in the pulse program, the HW_Trigger will cause the program to continue to the next instruction.

The HW_Reset line is pulled high by a resistor.  It can be used to halt the execution of a program by pulling it low (e.g., by shorting pins 5-6).  When the signal is pulled low during the execution of a program, the controller resets itself back to the beginning of the program.  Program execution can be resumed by either a software start command or by a hardware trigger.

Additional Output Signals – Connector DB-25 and Header J12.

The internal Output Clock signal is available to the outside world, as it is tied to the DB25 connector, pin #1.  It is the clock signal used to latch patterns in the output buffers.  This clock has been configured to have a relatively slow slew rate so as to avoid noise problems on a transmission line.  This clock is not a 50% duty cycle clock.  The width of the high part of the signal is one system clock period.  For a 100 MHz board, the signal is at a logical one for 10 ns at every transition to a new pulse program line.

System Reset – Header J12 pin #8 is used to indicate (when low) to the external world that the uPC controller is in a reset state.  It can be used in larger systems to monitor the state of the PulseBlaster™ design.

A signal that is similar to System Reset is the Running signal, header J12 pin #10.  It is driven high when the uPC is executing code.  It is taken low when the uPC has entered either a reset of idle state.

The Stopped signal, header J12 pin #7, is the last signal used to indicate the state of the uPC.  Stopped is asserted when the uPC has encountered the stop command while normally executing code.  This signal informs the external world that the uPC has successfully executed its program and has halted operation.

Header and Signal Locations

The location of the relevant headers and connectors on the PulseBlasterDDS™ board is presented in Figure 3.

SMA4   J4        SMA3 	SMA2   J_Power2  J_Power1

J10(DB-25)

Figure 3.  PulseBlasterDDS™ Board – header/connector locations.

##### J12

JTrigger

4. ISA Readback Feature

PulseBlaster and PulseBlasterDDS products have the ability to provide status information to the user through a group of registers that are readable through the ISA Bus.  The registers provide information on the status of the PulseBlaster series device as well as the firmware version number.

The register numbers provided below are offsets from the ISA port base address assigned to the PulseBlaster series card.  All registers are 8 bits and are read only.  If the users attempts to write to the port address associated with one of the registers, it will overwrite data in one of the IBC (ISA Bus Controller) control registers.

Only Register 0 and Register 1 are guaranteed to remain the same in future releases.

Register Definitions

Register 0 – Control Signal Register

Bit 7: Reserved (Testing := Memory_Mode)

Bit 6: IBC_Error – indicates if there has been an error programming the IBC

Bit 5: Programming_Finished – indicates that the PulseBlaster series device has been    programmed and is armed.  The next trigger (either Hardware or Software) will start execution of the pulse programmer code.

Bit 4: Idle -

Bit 3: Waiting – The PulseBlaster series device has encountered a WAIT Op Code and is currently waiting for the next trigger (either Hardware or Software) to resume operation.

Bit 2: Running – indicates that the PulseBlaster series device is current executing a program.

Bit 1: Reset – the PulseBlaster series device is in a RESET state and must be reprogrammed before code execution can begin again.

Bit 0: Stop – indicates that the PulseBlaster series device has encountered a STOP Op Code during program execution and has entered a stopped state.

Register 1 – Version Control Register

Bits[7..4]: Product Type Number

PulseBlaster (external memory) = 0

PulseBlaster (internal memory) = 1

PulseBlasterDDS (PP only external memory) = 2

PulseBlasterDDS (PP only internal memory) = 3

PulseBlasterDDS (single channel DDS) = 4

PulseBlasterDDS (dual channel integrated DDS) = 5

PulseBlasterDDS (2 integrated, 2 external DDS) = 6

PulseBlasterDDS (2 AWG, 2 DDS) = 7

Bits[3..0]: Release Version Number

Register 2 – Pulse Programmer Core Address LSB’s

Bits[7..0] : PP Core Current_Address[7..0]

Register 3 – Pulse Programmer Core Address MSB’s

Bits[7..0]: PP Core Current_Address[15..8]

Register 4 – PP Core Delay Counter Value (LSB)

Bits[7..0]: PP Core Delay_Count[7..0]

Register 5 - PP Core Delay Counter Value (MSB)

Bits[7..0]: PP Core Delay_Count[31..24]

Register 6

Bits[7]: PP Core Loop_Done

Bits[6..4]: PP Core Control_Mux[2..0]

Bits[3..0]: PP Core Control_Codes[3..0]

Register 7

Bits[7..0]: PP Core Branch_Address[7..0]

Section III: Programming Notes: Using C Functions (Provided)

There are two methods of programming the PulseBlasterDDS.  You can either use the provided C functions to create the program for you by putting your pulse program in the provided example source C file and recompiling it, or you can generate the appropriate outputs directly to the device using the “Write Byte” commands, as appropriate for your operating system and/or application software environment.  Please consult Appendix II for more information on using the direct method.

1.  Provided C Functions: Description

A series of functions has been written to make writing pulse programs easier.  In order to use these functions, the header file (pbdfuncs.h) and source file (pbdfuncs.c) must be in the working directory of your C compiler.  You can obtain these files from http://www.pulseblaster.com/CD/PulseBlasterDDS/ISA/

void set_ISA_address(int address);

Used to set the base address of the ISA card.  The default value is 0x340.  You only need to call this function if you are using an ISA card and your card’s base address is not 0x340.

void set_clock(double clock_freq);

Used to set the clock frequency of the system.  The variable clock_frequency is specified in MHz.  The default value is 50MHz.  You only need to call this function if you are not using a –50 system.

int start_programming(int device);

Used to initialize the system to receive programming information.  It accepts a parameter referencing the target for the instructions.  Valid values for device are PULSE_PROGRAM, FREQ_REGS, TX_PHASE_REGS (or equivalent PHASE_REGS), and RX_PHASE_REGS.  It returns a 0 on success or a negative number on an error.

int set_freq(double freq);

Used to set the values in the frequency registers.  Should only be called after start_programming(FREQ_REGS) has been called.  Registers are programmed one at a time, starting at 0 and incrementing each time this function is called.  It accepts the value for the frequency register in MHz.  It returns a 0 on success or a negative number on an error.

int set_phase(double phase);

Used to set the values in the phase registers.  Should only be called after start_programming(int device) has been called with a phase register target.  Registers are programmed one at a time, starting at 0 and incrementing each time this function is called.  It accepts the value for the phase register in degrees.  It returns a 0 on success or a negative number on an error.

int pb_inst(int freq, int tx_phase, int tx_output_enable, int rx_phase, int rx_output_enable, int flags, int inst, int inst_data, double length);

Used to send one instruction of the pulse program.  Should only be called after start_programming(PULSE_PROGRAM) has been called.  It returns a negative number on an error, or the instruction number upon success.  If the function returns –99, an invalid parameter was passed to the function.  Instructions are numbered starting at 0.

int freq – Number of the frequency register to be used.  Valid range is from 0 to 15.

int tx_phase – Number of the phase register to be used on the TX output channel.  Valid range is from 0 to 15.

int tx_output_enable – Determines whether analog output channel TX is generating a sinusoid or is at ground.  Valid values are TX_ANALOG_ON(or equivalent ANALOG_ON) and TX_ANALOG_OFF(or equivalent ANALOG_OFF).

int rx_phase – Number of the phase register to be used on the RX output channel.  Valid range is from 0 to 15.

int rx_output_enable – Determines whether analog output channel RX is generating a sinusoid or is at ground.  Valid values are RX_ANALOG_ON(or equivalent ANALOG_ON) and RX_ANALOG_OFF(or equivalent ANALOG_OFF).

int flags – determines state of each TTL output bit.  Valid values are 0x0 to 0x3FF.  For example, 0x010 would correspond to bit 7 being on, and all other bits being off.

int inst – determines which type of instruction is to be executed.  Valid instructions are included in table 2.

int inst_data – data to be used with the previous inst field.  The meaning of this is described in table 2.

double length – duration of this pulse program instruction, specified in ns.

This function has been overloaded to accommodate TTL-only programs or single-RF-channel programs.  When using a shorter version of the function, any RF channels not specified have their outputs set to ground.  The single channel RF function uses the TX output channel.  The overloaded forms follow:

Single Channel:

int pb_inst(int freq, int phase, int rf_output_enable, int flags, int inst, int inst_data, int length);

TTL Only:

int pb_inst(int flags, int inst, int inst_data, int length);

int stop_programming();

Used to tell that programming the system is complete.  Program execution cannot start until this command is received.  It returns a 0 on success or a negative number on an error.

int start_pb();

Once system has been programmed, this instruction will start execution of pulse program.  It returns a 0 on success or a negative number on an error.

int stop_pb();

Stops output of system.  Analog output will return to ground, and TTL outputs will remain in the state they were in when stop command was received.  It returns a 0 on success or a negative number on an error.

| Inst | inst_data | Meaning |
| --- | --- | --- |
| CONTINUE | Ignored | Program execution continues to next instruction |
| STOP | Ignored | Stop execution of program (*Note all TTL values remain from previous instruction, and analog outputs turn off) |
| LOOP | Number of desired loops. This value must be greater than or equal to 1. | Specify beginning of a loop. Execution continues to next instruction. Data used to specify number of loops |
| END_LOOP | Address of beginning of loop | Specify end of a loop. Execution returns to begging of loop and decrements loop counter. |
| JSR | Address of first subroutine instruction | Program execution jumps to beginning of a subroutine |
| RTS | Ignored | Program execution returns to instruction after JSR was called |
| BRANCH | Address of next instruction | Program execution continues at specified instruction |
| LONG_DELAY | Number of desired loops. This value must be greater than or equal to 2. | For long interval instructions. Data field specifies a multiplier of the delay field. Execution continues to next instruction |
| WAIT | Ignored | Program execution stops and waits for software or hardware trigger. Execution continues to next instruction after receipt of trigger. |

Table 4: Instruction Types and meaning of associated data.

2. Example Use of C Functions

The following example is a complete pulse program that can be executed on the PulseBlasterDDS™ board.  The program, upon executing, will load and initiate the PulseBlaterDDS™ board to start generating a sequence of 10 us RF pulses on Tx channel every 1 ms. The Rx channel will be turned ON during the 1 ms gap between the Tx pulses and OFF during the Tx pulse. The phase of the Rx signal will be phase shifted by 180 degrees with respect to the Rx phase.  The generated rf frequencies will be 1.012 MHz.  While executing, the TTL outputs will toggle between the logical “one” during the 10 us intervals and logical “zero” during the 1 ms intervals.

// Filename: Example1.cpp

//Sample Program for PulseBlasterDDS

#include <stdio.h>

#include "pbdfuncs.cpp"

#include "pbdfuncs.h"

void main()

{

int start;

set_clock(100); // 100 MHz reference clock frequency

start_programming(FREQ_REGS);

set_freq(1.012); // 1.012 MHz in 0th frequency register

start_programming(TX_PHASE_REGS);

set_phase(0); // Tx phase value set to 0 degrees (in 0th phase register)

start_programming(RX_PHASE_REGS);

set_phase(180); // Rx set to 180 degrees out of phase

start_programming(PULSE_PROGRAM);

// Instruction 0 - Continue to next instruction in 10.0 us

start = pb_inst(0,0,TX_ANALOG_ON,0,RX_ANALOG_ON,0x3FF,CONTINUE,0,10.0*us);

// Instruction 1 - Branch to "start" (Instruction 0) in 1 ms

pb_inst(0,0,TX_ANALOG_OFF,0,RX_ANALOG_ON,0x0,BRANCH,start,1*ms);

stop_programming();

start_pb(); //the pulse program  will start

}

A more complex pulse program example is provided in the Appendix I.

Appendix I: Programming Notes: Sample C Program

// Filename: Example2.cpp

// SpinCore Technologies, Inc.

// August 2002

// http://www.spincore.com

#include <stdio.h>

#include "pbdfuncs.cpp"

#include "pbdfuncs.h"

void main()

{

// Set clock frequency

set_clock(50);

// Start programming the frequency registers

start_programming(FREQ_REGS);

// Program the registers in order from 0 to 15

// Valid units are MHz, kHz, and Hz - default is MHz

set_freq(1*MHz); // Set register 0

set_freq(2*MHz); // Set register 1

set_freq(3*MHz); // Set register 2

set_freq(4*MHz); // Set register 3

set_freq(5*MHz); // Set register 4

set_freq(6*MHz); // Set register 5

set_freq(7*MHz); // Set register 6

set_freq(8*MHz); // Set register 7

set_freq(9*MHz); // Set register 8

set_freq(10*MHz); // Set register 9

set_freq(11*MHz); // Set register 10

set_freq(12*MHz); // Set register 11

set_freq(13*MHz); // Set register 12

set_freq(14*MHz); // Set register 13

set_freq(15*MHz); // Set register 14

set_freq(16*MHz); // Set register 15

// Start programming the TX channel phase registers

start_programming(TX_PHASE_REGS);

// Program the registers in order from 0 to 15, Units are in degrees

set_phase(0); // Set register 0

set_phase(22.5); // Set register 1

set_phase(45); // Set register 2

set_phase(67.5); // Set register 3

set_phase(90); // Set register 4

set_phase(112.5); // Set register 5

set_phase(135); // Set register 6

set_phase(157.5); // Set register 7

set_phase(180); // Set register 8

set_phase(202.5); // Set register 9

set_phase(225); // Set register 10

set_phase(247.5); // Set register 11

set_phase(270); // Set register 12

set_phase(292.5); // Set register 13

set_phase(315); // Set register 14

set_phase(337.5); // Set register 15

// Start programming the RX channel phase registers

start_programming(RX_PHASE_REGS);

// Program the registers in order from 0 to 15, Units are in degrees

set_phase(0); // Set register 0

set_phase(22.5); // Set register 1

set_phase(45); // Set register 2

set_phase(67.5); // Set register 3

set_phase(90); // Set register 4

set_phase(112.5); // Set register 5

set_phase(135); // Set register 6

set_phase(157.5); // Set register 7

set_phase(180); // Set register 8

set_phase(202.5); // Set register 9

set_phase(225); // Set register 10

set_phase(247.5); // Set register 11

set_phase(270); // Set register 12

set_phase(292.5); // Set register 13

set_phase(315); // Set register 14

set_phase(337.5); // Set register 15

//Begin pulse program

start_programming(PULSE_PROGRAM);

int start, loop, sub; // define instruction labels

sub = 5; // Since we are going to jump forward in our program, we need to

// define this variable by hand.  Instructions start at 0 and count up

//Instruction format

//int pb_inst(int freq, int tx_phase, int tx_output_enable, int rx_phase, int rx_output_enable, int flags, int inst, int inst_data, int length)

// Instruction 0 - Jump to Subroutine at Instruction 5 in 1us

start =	pb_inst(0,0,TX_ANALOG_OFF,0,RX_ANALOG_OFF, 0x3FF,JSR,sub,1*us);

// Instruction 1 - Beginning of Loop, loop 3 times, Continue to next instruction in 1us

loop =	pb_inst(0,0, TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,LOOP,3,1*us);

// Instruction 2 - End of Loop.  Return to beginning of loop or continue to next instruction in 1us

pb_inst(0,0, TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,END_LOOP,loop,1*us);

// Instruction 3 - Stay here for (5*1us) then continue to Instruction 4

pb_inst(0,0, TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,LONG_DELAY,5,1*us);

// Instruction 4 - Branch to "start" (Instruction 0) in 1us

pb_inst(0,0, TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,BRANCH,start,1*us);

// Instruction 5 - Continue to next instruction in 2us

pb_inst(0,0, TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,CONTINUE,0,2*us);

// Instruction 6 - Return from Subroutine to Instruction 1 in 2us

pb_inst(0,0, TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,RTS,0,2*us);

// End of programming registers and pulse program

stop_programming();

// Start execution of the pulse program

start_pb();

}

Appendix II: Programming Notes: Writing Directly to I/O Ports

PulseBlasterDDS™ boards can be programmed by directly writing bytes to specific I/O ports of a host computer. Programming of the PulseBlasterDDS™ board for operation involves a minimum of four steps.  The steps are as follows:

- Send LOAD NUMBER OF BYTES PER WORD instruction.

- Send SELECT PERIPHERAL DEVICE instruction.

- Send CLEAR ADDRESS COUNTER instruction.

3.A. (Optional) loading data to memory.

- Send PROGRAMMING FINISHED instruction.

If these four commands are not sent from a PC, the PulseBlasterDDS™ board will not run as desired.  All four instructions are required as an attempt to ensure that the device has been programmed before it can be armed.  The first time the board is used, the loading of the memory with data has to be performed between steps three and four, step 3.A above.  Upon reset, all four instructions must be executed to restart the device again.

Formulas to calculate the byte values to program the DDS frequency, phase, and timing registers are provided in the example programming sequence later in this section.  Additional explanation of the programming registers follows.

WE for Peripherals: This register is used to select the peripheral that is to be programmed.  The value of this register that is used to select program memory is always zero and this is the default value for the register.  A complete listing of the values and the associated hardware that can be programmed when appropriately set.

Table 5: Peripheral List

CLEAR ADDRESS COUNTER: The Address Counter is used to manufacture the memory address.  The Address Counter is not loadable; it can only be cleared and started at zero.  It is not possible to load a particular section of memory.  All loads must start from either the beginning of memory, or wherever the Address Counter left off.

Flag Initialization Strobe:  The output flags of the PulseBlasterDDS can be programmed while the device is in a reset state.  This is useful to initialize flags after powering-up and to reset flags to a known state if a program must be aborted.  Writing to the Flag Initialization Strobe register will toggle the line used to clock data into the output latches.  The example below provides more information on how to use the Flag Initialization Strobe to program the output flags while the PulseBlasterDDS is in a reset state.

LOAD_MEMORY: This instruction is used to specify data that should be used to program the memory used by the device.  Since the ISA data is taken only one byte at a time, the IBC must reconstruct the data word to be programmed.  The data word is reconstructed in the IBC most significant byte first.

PROGRAMMING FINISHED: This instruction enables the pattern generator of the PulseBlaster™.  This instruction prevents the pattern generator from accepting a hardware trigger or software start command before the device has been programmed.  Once the design has been programmed, the PROGRAMMING FINISIHED command must be sent to arm the device for operation.  After the pattern generator has been armed, any hardware trigger or software start command will cause the system to start operation. The PulseBlaster™ can be reset by issuing the DEVICE_RESET command.  This will internally clear the PROGRAMMING FINISHED instruction and prevent the pattern generator from operating again until the IBC has been re-initialized.

Example output sequence

The following is an example of the output sequence to program the PulseBlasterDDS board.  Explanations are included in brackets in the middle of the code.

[	Initialization:	]

Output "0x00" to port base + 0 (Issue device reset)

Output "0x04" to port base + 2 (Select number of bytes per word)

Output "0xFF" to port base + 3 (Select device to program (Flag initial values))

Output "0x00" to port base + 4 (Reset address counter)

[	Set initial flag values	]

[	Values for this example are "0x00aaf0f0"	]

Output "0x00" to port base + 6 (Data transfer)

Output "0xAA" to port base + 6 (Data transfer)

Output "0xF0" to port base + 6 (Data transfer)

Output "0xF0" to port base + 6 (Data transfer)

Output "0x00" to port base + 5 (Clock data into external buffer)

Output "0x00" to port base + 5 (Return clock signal to low)

[	Set up DDS frequency registers	]

Output "0x00" to port base + 0 (Issue device reset)

Output "0x04" to port base + 2 (Select number of bytes per word)

Output "0x01" to port base + 3 (Select device to program (DDS Frequency Registers))

Output "0x00" to port base + 4 (Reset address counter)

[	DDS Register Values	]

[	Reg0 = 051EB852 (1 MHz)	]

[	Reg1 = 0A3D70A4 (2 MHz)	]

[	Reg2 = 0F5C28F6 (3 MHz)	]

[	Reg3 = 147AE148 (4 MHz)	]

[	Formula for finding these values:	]

[	REG0 = DESIRED_FREQUENCY * 232 / PBDDS_CLOCK			]

[		= 1 MHz * 232 / 50 MHz = 858993459.2 = 0x051EB852	]

Output "0x05" to port base + 6 (Data Transfer - Byte 3 of Reg0)

Output "0x1E" to port base + 6 (Data Transfer - Byte 2 of Reg0)

Output "0xB8" to port base + 6 (Data Transfer - Byte 1 of Reg0)

Output "0x52" to port base + 6 (Data Transfer - Byte 0 of Reg0)

Output "0x05" to port base + 6 (Data Transfer - Byte 3 of Reg1)

Output "0x3D" to port base + 6 (Data Transfer - Byte 2 of Reg1)

Output "0x70" to port base + 6 (Data Transfer - Byte 1 of Reg1)

Output "0xA4" to port base + 6 (Data Transfer - Byte 0 of Reg1)

Output "0x0F" to port base + 6 (Data Transfer - Byte 3 of Reg2)

Output "0x5C" to port base + 6 (Data Transfer - Byte 2 of Reg2)

Output "0x28" to port base + 6 (Data Transfer - Byte 1 of Reg2)

Output "0xF6" to port base + 6 (Data Transfer - Byte 0 of Reg2)

Output "0x14" to port base + 6 (Data Transfer - Byte 3 of Reg3)

Output "0x7A" to port base + 6 (Data Transfer - Byte 2 of Reg3)

Output "0xE1" to port base + 6 (Data Transfer - Byte 1 of Reg3)

Output "0x48" to port base + 6 (Data Transfer - Byte 0 of Reg3)

[	Set up TX phase registers	]

Output "0x00" to port base + 0 (Issue device reset)

Output "0x04" to port base + 2 (Select number of bytes per word)

Output "0x02" to port base + 3 (Select device to program (TX Phase Registers))

Output "0x00" to port base + 4 (Reset address counter)

[	TX Phase Register Values	]

[	Reg0 = 093E93E9 (13 degrees)	]

[	Reg1 = 40000000 (90 degrees)	]

[	.					]

[	.					]

[	Reg15 = E0000000(315 degrees)	]

[	Formula for finding these values:	]

[	REG0 = DESIRED_PHASE * 232 / 360				]

[		= 13 * 232 / 360 = 155096041.24  0x093E93E9	]

[	* Note that only the top 3 hex digits (0x093) are actually used	]

[	in programming the phase register.  The other digits must still	]

[	be sent so that the dds word remains 32 bits long.			]

Output "0x09" to port base + 6 (Data Transfer - Byte 3 of Reg0)

Output "0x3E" to port base + 6 (Data Transfer - Byte 2 of Reg0)

Output "0x93" to port base + 6 (Data Transfer - Byte 1 of Reg0)

Output "0xE9" to port base + 6 (Data Transfer - Byte 0 of Reg0)

Output "0x40" to port base + 6 (Data Transfer - Byte 3 of Reg1)

Output "0x00" to port base + 6 (Data Transfer - Byte 2 of Reg1)

Output "0x00" to port base + 6 (Data Transfer - Byte 1 of Reg1)

Output "0x00" to port base + 6 (Data Transfer - Byte 0 of Reg1)

(…)

Output "0xE0" to port base + 6 (Data Transfer - Byte 3 of Reg15)

Output "0x00" to port base + 6 (Data Transfer - Byte 2 of Reg15)

Output "0x00" to port base + 6 (Data Transfer - Byte 1 of Reg15)

Output "0x00" to port base + 6 (Data Transfer - Byte 0 of Reg15)

[	Pulse Program Setup	]

Output "0x00" to port base + 0 (Issue Device Reset)

Output "0x0A" to port base + 2 (Select number of bytes per word)

Output "0x00" to port base + 3 (Select device to program (RAM))

Output "0x00" to port base + 4 (Reset address counter)

Output "0x18" to port base + 6 (Byte 9 of first instruction)

Output "0xFF" to port base + 6 (Byte 8 of first instruction)

Output "0xFF" to port base + 6 (Byte 7 of first instruction)

Output "0x00" to port base + 6 (Byte 6 of first instruction)

Output "0x00" to port base + 6 (Byte 5 of first instruction)

Output "0x00" to port base + 6 (Byte 4 of first instruction)

Output "0x00" to port base + 6 (Byte 3 of first instruction)

Output "0x00" to port base + 6 (Byte 2 of first instruction)

Output "0x00" to port base + 6 (Byte 1 of first instruction)

Output "0x07" to port base + 6 (Byte 0 of first instruction)

Output "0xXX" to port base + 6 (Byte 9 of second instruction)

[	Continue this process for all instructions.  The	]

[	explanation of how to create the 80 bit instruction]

[	words is included below.  When finished with all	]

[	instructions, continue with the sequence below.		]

Output "0x00" to port base + 7 (Programming Finished)

[	Only execute the following command when you are ready	]

[	for the program to start running.				]

Output "0x00" to port base + 1 (Start pulse program)