PulseBlasterDDS

[image: image3.png]DDS Core

Tomerically
Controlled
Oseillator

Reference
Oscillator

Outputs

Output and Control Register

2002 SpinCore Technologies, Inc.

User Control -
http://www spincore.com




[image: image4.png]DDS Core

Tomerically
Controlled
Oseillator

Reference
Oscillator

Outputs

Output and Control Register

2002 SpinCore Technologies, Inc.

User Control -
http://www spincore.com




 
PulseBlasterDDS™
PCI Board Rev. 02

Owner’s Manual 


[image: image5.png]O
SinCore

Technologies, Inc.



Models:

[image: image6.jpg]


PBDDS-100-PCI

PBDDS-50-PCI

SpinCore Technologies, Inc.

http://www.spincore.com 

[image: image7.png]Found Ney

Hardware

izard

‘Welcome to the Found New

Hardware Wizard

This wizard helps you nstal software for
PulseBlasteDDS PEDUIPC

1f your hardware came with an installation CD
& or Hloppy disk. inser it now.

What do you want the wizaid o do?

Instal the saftware automaticaly (Recommended)
il o & 161 specie osation (Advanced]

Clck Net o contiue.

<Back





Congratulations and thank you for choosing a design from SpinCore Technologies, Inc.


We appreciate your business!

At SpinCore we try to fully support the needs of our customers.  If you are in need of assistance, please contact us and we will strive to provide the necessary support.
[image: image8.png]Found New Hardware Wizard

Please choose your search and installation options.

(@ Search for the best driver in these locations.

Use the check boses below o it o expand the default search, which includss local
paths and removable medis. The best diver found wil be nstalled

[ Search removable media (floppy. CD-AOM...]
il 1 geaion e seaicy

c\pbest

O Don't ssarch. | will choose the diiver to nstall

Chaase this optian to selec the device civer fiom a s, Windows does not quarantee that
the dhiver you choose wil be the best malch far your hardware.





© 2000-2002 SpinCore Technologies, Inc. All rights reserved.
SpinCore Technologies, Inc. reserves the right to make changes to the product(s) or information herein without notice. PulseBlasterDDS™, PulseBlaster™, SpinCore, and the SpinCore Technologies, Inc. logos are trademarks of SpinCore Technologies, Inc. All other trademarks are the property of their respective owners.

SpinCore Technologies, Inc. makes every effort to verify the correct operation of the equipment.  This equipment version is not intended for use in a system in which the failure of a SpinCore device will threaten the safety of equipment or person(s).
Table of Contents 

3I. Introduction

Product Overview
3
Board Architecture
4
Block Diagram
4
Output signals
4
Timing characteristics
5
Instruction set
5
External triggering
5
Versatile DAC
5
Summary
6
Specifications
6
DDS Specifications (Model DDS-100)
6
TTL Specifications
6
Common Parameters (DDS and TTL Specifications)
6
Pulse Program Control Flow (Common)
6
II. Installation
7
Installing the PulseBlasterDDS Driver
7
For Windows XP
7
III. Programming the PulseBlasterDDS
9
Instruction Set Architecture
9
Machine-Word Definition
9
Breakdown of 80-bit Instruction Word
9
Using C Functions to Program the PulseBlasterDDS
11
Example Use of C Functions
13
IV. Connecting to the PulseBlasterDDS Board
14
Connector Information
14
SMA Connector labeled “Wave_Out”
14
DB-25 - TTL Output Signal Bits
14
Header JP100
14
SMA Connector labeled “Ext_Trig”
15
SMA Connector labeled “Ext_Clk”
15
Appendix I: Sample C program
16
Example Program
16
Appendix II: Programming the PulseBlasterDDS Using Direct Outputs
19
Using DLL Functions to Send Instructions
19
Building Instructions Using the DLL Functions
19
Programming Information
19
Example Program
21
Contact Information
24



I. Introduction

Product Overview 

The PulseBlasterDDS( series of Intelligent Pattern and Waveform Generation boards from SpinCore Technologies, Inc., couples SpinCore’s unique Intelligent Pattern Generation processor core, dubbed PulseBlaster(, with Direct Digital Synthesis (DDS) for use in system control and waveform generation.

The PulseBlaster’s state-of-the-art timing processor core provides all the necessary timing control signals required for overall system control and waveform synchronization.   By adding DDS features, PulseBlasterDDS can now provide not only digital (TTL) but also analog output signals, meeting high-performance and high-precision complex excitation/stimuli needs of demanding users. 

PulseBlasterDDS provides users the ability to control their systems through the generation of fully synchronized (digital and analog) excitation waveforms from a small form factor PC board, providing users a compelling price/performance proposition unmatched by any other device on the market today.  Figure 1 presents sample capabilities of the board.

[image: image1.jpg]




Figure 1: Sample PulseBlasterDDS output capabilities 

Board Architecture 

Block Diagram 

Figure 2 presents the general architecture of the PulseBlasterDDS system.  The two major building blocks are the DDS Core and the Pulse Programming and Timing Processor Core (PP Core). The DDS Core contains a numerically controlled oscillator and has 16 programmable frequency registers that are under the pulse program control.  Prior to gating, the DDS signal can be phase offset by one of 16 programmable phase registers.  An optional modulator can be placed in the path of the output channel.  The PP Core controls the timing of the gating pulses and provides the necessary control signals for frequency and phase registers.  The DDS and PP cores have been integrated onto a single silicon chip.  High performance DAC chips and high current output amplifiers complement the design.  User control to the system is provided through the host-programming interface over the PCI bus.
[image: image9.png]Files Needed|

., Theie windivrsys’on SpinCore Technologis, Inc.
S Installation Disk is needed.

Type the path where thefile islocated, and then cick
oK.

oy fies fom:

c\pbest





Figure 2: PulseBlasterDDS board architecture

Output signals

The PulseBlasterDDS comes with one analog output channel that is configured to output radio-frequency (RF/IF) waveforms, and 10 digital output signal lines.  The frequency and phase of the RF waveform generated by the DDS is under the complete control of the user and are specified through software programming.  PulseBlasterDDS also provides the ability to gate the output of the DDS channel allowing for independent pulsed RF operation.  With digital sampling 


rate of 100 MHz (max. reference clock frequency), the maximum theoretical output frequency is 50 MHz (the Nyquist Theorem).  However with only two samples per period the usefulness of such outputs would be seriously limited.  In practice, at least four samples per period would be required to faithfully represent a signal.  Thus, at 100 MHz reference clock oscillator, analog signals up to approx. 25-30 MHz can be generated.  The analog output signal is available on an on-board SMA connector.  The output impedance of the analog signal is 50-ohms.

The 10 individually controlled digital (TTL/CMOS) output bits are capable of delivering (25 mA per bit and have an output voltage of 3.3V. These signals are available on the PC bracket-mounted DB-25 connector. 

Timing characteristics

PulseBlasterDDS’s timing controller can accept either an internal (on-board) crystal oscillator or an external frequency source of up to 100 MHz. The innovative architecture of the timing controller allows the processing of either simple timing instructions (delays of up to 232 = 4,294,967,296 clock cycles), or double-length timing instructions (up to 252 clock cycles long – nearly 2 years with a 100 MHz clock!).  Regardless of the type of timing instruction, the timing resolution remains constant for any delay – just one clock period (e.g., 10 ns for a 100 MHz clock). 

The core-timing controller has a very short minimum delay cycle – only five clock periods for internal memory (512 words) models.  This translates to a 50 ns pulse/delay/update with a 100 MHz clock. The external memory models (up to 32k words) have a nine clock-period minimum instruction cycle.

Instruction set

PulseBlasterDDS’ design features a set of commands for highly flexible program flow control. The micro-programmed controller allows for programs to include branches, subroutines, and loops at up to 8 nested levels – all this to assist the user in creating dense pulse programs that cycle through repetitious events, especially useful in numerous multidimensional spectroscopy and imaging applications. 

External triggering 

PulseBlasterDDS can be triggered and/or reset externally via dedicated hardware lines.  The two separate lines combine the convenience of triggering (e.g., in cardiac gating) with the safety of the "stop/reset" line.  The required control signals are “active low” (or short to ground).

Versatile DAC

PulseBlasterDDS PCI board rev. 02 is equipped with a versatile DAC that can operate in one of two modes: standard or PLL.  The PLL mode doubles the output rate at the cost of introducing a small phase noise.   Please contact sales@spincore.com to request designs specifically using either of these modes.
Summary

PulseBlasterDDS is a versatile, high-performance pulse/pattern TTL and RF/IF generator operating at speeds of up to 100 MHz and capable of generating pulses/delays/intervals ranging from 50 ns to over 2 years per instruction. It can accommodate pulse programs with highly flexible control commands of up to 32k program words. Its high-current output logic bits are independently controlled with a voltage of 3.3 V .  The output impedance of the analog channel is 50-ohms. 

Specifications

DDS Specifications (Model DDS-100)

· 100 MHz reference clock oscillator (other frequencies available upon request) 

· 0.047 Hz frequency resolution (32 bits)  

· 16 loadable frequency registers for agile frequency modulation/switching/selection (32 bits each) 

· 16 loadable phase-offset registers for agile phase modulation/switching/selection (12 bits each) 

· 0.09( phase resolution (12 bits)

· 40 ns phase switching latency 

· 40 ns frequency switching latency (phase continuous) 

· 10 dBm RF output power 

· 50 ohm output impedance 

· SMA connectors

· 60 MHz 3dB bandwidth

· RF Output capable of outputting DC at programmed output level (using phase offset)

TTL Specifications

· 10 individually controlled digital output lines (TTL levels)

· variable pulses/delays for every TTL line

· 25 mA output current per TTL line

· output lines can be combined to increase the max. output current

Common Parameters (DDS and TTL Specifications)

· 90 ns shortest pulse/interval

· 2 years longest pulse/interval

· 10 ns pulse/interval resolution

· RF and TTL pulses are synchronized

· 32k max. memory space

· external  triggering and reset – TTL levels

Pulse Program Control Flow (Common)

· loops, nested 8 levels deep

· 20 bit loop counters (max. 1,048,576 repetitions)

· subroutines, nested 8 levels deep

· wait for trigger - 80 ns latency, adjustable to 2 years in duration

· 5 MHz max. re-triggering frequency

II. Installation

Installing the PulseBlasterDDS Driver

***  The following section pertains to the Windows 98 Operating System.  If you require instructions for other operating systems, including Windows 2000 or Windows XP, please contact support@spincore.com.


1. Go to http://www.pulseblaster.com/CD/PulseBlasterDDS/PCI/PBD03PC and download PBD03PC.zip.


2. Unzip the files to their own directory.


3. Turn off your computer.


4. Insert the PulseBlasterDDS board into an empty PCI slot.


5. Turn on your computer.


For Windows XP

6.  After booting, the “Found New Hardware Wizard” should appear.  Choose “Install from a list or specific location” and click Next

[image: image10.png]Found Ney

Hardware Wizard

Completing the Found New
\Q\ Hardware Wizard
N

The wizard has frished instaling the saftware fo.

\> PulseBlasteDDS PEDUIPC

Clck Firsh to close the wizard.

<ok =





7.  Choose “Include this location in the search” and browse to the directory you unzipped the drivers to.  Click next. 


8.  While windows installs the driver, a “Files Needed” dialog may pop up.  Choose the directory you unzipped the drivers to, and click ok.


9.  When finished, you should see this window.


You are now ready to control the PulseBlasterDDS board
3. Run the included “PBD_Test.exe”.  


The board should now output a 3.125MHz sine wave on the SMA connector labeled “Wave_Out” NOTE: 100MHz Clock Models will output a 6.250MHz signal.  

The PulseBlasterDDS board is now ready for use!
III. Programming the PulseBlasterDDS

Instruction Set Architecture

Machine-Word Definition

The PulseBlaster pulse timing and control processor implements an 80-bit wide Very Long Instruction Word (VLIW) architecture. The VLIW memory words have specific bits/fields dedicated to specific purposes, and every word should be viewed as a single instruction of the micro-controller. The maximum number of instructions that can be loaded to on-board memory is 32k. The execution time of instructions can be varied and is under (self) control by one of the fields of the instruction word – the shortest being five clock cycles (for 512 memory-word models) and the longest being 2^52 clock cycles.  All instructions have the same format and bit length, and all bit fields have to be filled.  Figure 3 shows the fields and bit definitions of the 80-bit instruction word.  


Bit Definitions for the 80-bit Instruction Word (VLIW)
       Output/Control Word   |    Data Field    |     OP Code    |    Delay Count 

  (24 bits)             (20 bits)          (4 bits)         (32 bits)   

Figure 3: Bit definitions of the 80-bit instruction/memory word

Breakdown of 80-bit Instruction Word

The 80-bit VLIW is broken up into 4 sections

1. Output Pattern and Control Word - 24 bits 

2. Data Field - 20 bits

3. OP Code - 4 bits

4. Delay Count - 32 bits
Output Pattern and Control Word

Please refer to Table 1 for output pattern and control bit assignments of the 24-bit output/control word.

	Bit #
	
	Bit #
	Function

	23
	Selects Frequency Register (bit3)
	11
	Enables output of TX signal  (0 = on, 1 = off)

	22
	Selects Frequency Register (bit2)
	10
	Reserved

	21
	Selects Frequency Register (bit 1)
	9
	Output Connector DB25 pin 19

	20
	Selects Frequency Register (bit 0)
	8
	Output Connector DB25 pin 7

	19
	Selects Phase Register (bit 3)
	7
	Output Connector DB25 pin 8

	18
	Selects Phase Register (bit 2)
	6
	Output Connector DB25 pin 21

	17
	Selects Phase Register (bit 1)
	5
	Output Connector DB25 pin 22

	16
	Selects Phase Register (bit 0)
	4
	Output Connector DB25 pin 10

	15
	Reserved
	3
	Output Connector DB25 pin 11

	14
	Reserved
	2
	Output Connector DB25 pin 24

	13
	Reserved
	1
	Output Connector DB25 pin 25

	12
	Reserved
	0
	Output Connector DB25 pin 13


Table 1: Output Pattern and Control Word Bits

Data Field and Op Code

Please refer to Table 2 for information on the available operational codes (OpCode) and the associated data field functions (the data field's function is dependent on the Op Code)

	Op Code #
	Inst
	Inst_data
	Function

	0
	CONTINUE
	Ignored
	Program execution continues to next instruction

	1
	STOP
	Ignored
	Stop execution of program (*Note all TTL values remain from previous instruction, and analog outputs turn off)

	2
	LOOP
	Number of desired loops.  This value must be greater than or equal to 1.
	Specify beginning of a loop.  Execution continues to next instruction.  Data used to specify number of loops

	3
	END_LOOP
	Address of beginning of loop
	Specify end of a loop.  Execution returns to begging of loop and decrements loop counter.

	4
	JSR
	Address of first subroutine instruction
	Program execution jumps to beginning of a subroutine

	5
	RTS
	Ignored
	Program execution returns to instruction after JSR was called

	6
	BRANCH
	Address of next instruction
	Program execution continues at specified instruction

	7
	LONG_DELAY
	Number of desired loops.  This value must be greater than or equal to 2.
	For long interval instructions.  Data field specifies a multiplier of the delay field.  Execution continues to next instruction

	8
	WAIT
	Ignored
	Program execution stops and waits for software or hardware trigger.  Execution continues to next instruction after receipt of trigger. The latency is equal to the delay value entered in the WAIT instruction line plus a fixed delay of 6 clock cycles.


Table 2: Op Code and Data Field Description
Delay Count

The value of the Delay Count field (a 32-bit value) determines how long the current instruction should be executed.  The allowed minimum value of this field is 0x6 for the 32k memory models and 0x2 for the internal-memory models.  The timing controller has a fixed delay of three clock cycles and the value that one enters into the Delay Count field should account for this inherent delay.

Using C Functions to Program the PulseBlasterDDS

A series of functions have been written to control the board and facilitate the construction of pulse program instructions.  The functions also allow the programmer to set the DDS frequency and phase registers.

In order to use these functions, the DLL (pbd03pc.dll), the library file (pbd03pc.lib), the header files (pbd03pc.h and pbdfuncs.h), and source file (pbdfuncs.cpp) must be in the working directory of your C compiler
.

int pb_init();

Initializes PulseBlasterDDS board.  Needs to be called before calling any functions using the PulseBlasterDDS.  Returns a negative number on an error or 0 on success.

int pb_close();

Releases PulseBlasterDDS board.  Needs to be called as last command in pulse program.  Returns a negative number on an error or 0 on success.

void set_clock(double clock_freq);

Used to set the clock frequency of the board.  The variable clock_frequency is specified in MHz when no units are entered.  Valid units are MHz, kHz, and Hz.  The default clock value is 50MHz.  You only need to call this function if you are not using a –50 board.

int start_programming(int device);
Used to initialize the system to receive programming information.  It accepts a parameter referencing the target for the instructions.  Valid values for device are PULSE_PROGRAM, FREQ_REGS, and PHASE_REGS.  It returns a 0 on success or a negative number on an error.

int set_freq(double freq);
Used to set the values in the frequency registers.  Should only be called after start_programming(FREQ_REGS) has been called.  Registers are programmed one at a time, starting at 0 and incrementing each time this function is called.  It accepts the value for the frequency register with a default unit of MHz.  Valid units are MHz, kHz, Hz.  It returns a 0 on success or a negative number on an error.

int set_phase(double phase);
Used to set the values in the phase registers.  Should only be called after start_programming(PHASE_REGS) has been called.  Registers are programmed one at a time, starting at 0 and incrementing each time this function is called.  It accepts the value for the phase register in degrees.  It returns a 0 on success or a negative number on an error.

int pb_inst(int freq, int phase, int rf_output_enable, int flags,

int inst, int inst_data, double length);
Used to send one instruction of the pulse program.  Should only be called after start_programming(PULSE_PROGRAM) has been called.  It returns a negative number on an error, or the instruction number upon success.  If the function returns –99, an invalid parameter was passed to the function.  Instructions are numbered starting at 0.

int freq – Number of the frequency register to be used.  Valid range is from 0 to 16

int phase – Number of the phase register to be used.  Valid range is from 0 to 16

int rf_output_enable – Determines whether analog output is generating a sinusoid or is at ground.  Valid values are ANALOG_ON and ANALOG_OFF


int flags – determines state of each TTL output bit.  Valid values are 0x0 to 0x3FF.  For 
example, 0x010 would correspond to bit 5 being on, and all other bits being off.

int inst – determines which type of instruction is to be executed. Please see Table 2 for details.

int inst_data – data to be used with the previous inst field.  Please see Table 2 for details.

double length – duration of this pulse program instruction, specified in ns.

This function has been overloaded to accommodate TTL-only programs.  When using the shorter version of the function, the RF channel has its output set to ground. The overloaded form follow:

TTL Only:

int pb_inst(int flags, int inst, int inst_data, double length);

int stop_programming();
Used to tell that programming the board is complete.  Board execution cannot start until this command is received.  It returns a 0 on success or a negative number on an error.

int start_pb();
Once board has been programmed, this instruction will start execution of pulse program.  It returns a 0 on success or a negative number on an error.

int stop_pb();
Stops output of board.  Analog output will return to ground, and TTL outputs will remain in the state they were in when stop command was received.  It returns a 0 on success or a negative number on an error.

Example Use of C Functions


// Example1.cpp


//


// SpinCore Technologies, Inc.


// October 2002


// http://www.spincore.com


//


// The following program code uses C Functions from 'pbdfuncs' to


// generate and execute a pulse sequence on the PulseBlasterDDS 


// board.


// Be sure to include the DLL (pbd03pc.dll), the library file


// (pbd03pc.lib), the header files (pbd03pc.h and pbdfuncs.h), and 


// source file (pbdfuncs.cpp) in the working directory of your C


// compiler.

#include <stdio.h>

#include "pbdfuncs.h"

#include "pbdfuncs.cpp"

void main(void)


{


pb_init();  // Locates & Initializes the PulseBlasterDDS Board


int start;


set_clock(100);  // Set Because the Board Operates at 100MHz


start_programming(FREQ_REGS);  // Prepare the Board to Receive


set_freq(1.054);


 // Load Frequency Register 0


set_freq(2);


 // Load Frequency Register 1


start_programming(PULSE_PROGRAM); // Prepare the Board to Receive







    // Pulse Program Instructions


// Instruction 0 - Continue to instruction 1 in 2us


// DDS Output ON, Flags = 0x3FF, OPCODE = CONTINUE


start = pb_inst(0, 0, ANALOG_ON, 0x3FF, CONTINUE, 0, 2*us);


// Instruction 1 - Continue to instruction 2 in 0.1ms


// DDS Output OFF, Flags = 0x0, OPCODE = CONTINUE


pb_inst(0, 0, ANALOG_OFF, 0x0, CONTINUE, 0, 0.1*ms);


// Instruction 2 - Branch to "start" (Instruction 0) in 0.1s


// DDS Output OFF, Flags = 0x0, OPCODE = BRANCH, Target = start


pb_inst(0, 0, ANALOG_OFF, 0x0, BRANCH, start, 0.1*s);


stop_programming();

    // Finished Sending Instructions


start_pb();



    // Run the Program

pb_close();



    // Release Control of the 

    
    // PulseBlasterDDS board

}

A more complex program using C Functions is provided in Appendix I.
IV. Connecting to the PulseBlasterDDS Board

Connector Information

SMA Connector labeled “Wave_Out”

Outputs DDS signals generated by the user’s Program. 
DB-25 - TTL Output Signal Bits

Outputs TTL signals generated by the user’s Program.  Please consult the table below for bit assignments.

	Pin Assignments

	Pin#
	Bit#
	Pin#
	Bit#

	1
	GND
	14
	GND

	2
	Reserved
	15
	Reserved

	3
	GND
	16
	Reserved

	4
	Reserved
	17
	GND

	5
	Reserved
	18
	Reserved

	6
	GND
	19
	9

	7
	8
	20
	GND

	8
	7
	21
	6

	9
	GND
	22
	5

	10
	4
	23
	GND

	11
	3
	24
	2

	12
	GND
	25
	1

	13
	0
	
	


Table 3: Output bits and signals of the PulseBlasterDDS board
Header JP100


This is an input connector, for hardware triggering (HW_Trigger) and resetting (HW_Reset).

HW_Trigger is pulled high by default, and pin 1 is active (pin 2 = GND).  When a falling edge is detected (e.g., when shorting pins 1-2), it initiates code execution.   This trigger will also restart execution of a program from the beginning of the code if it is detected after the design has reached an idle state.  The idle state could have been created either by reaching the STOP Op Code of a program, or by the detection of the HW_Reset signal.  When the WAIT Op Code is used in the pulse program, the HW_Trigger will cause the program to continue to the next instruction.

HW_Reset is pulled high by default, and pin 3 is active (pin 4 = GND).  It can be used to halt the execution of a program by pulling it low (e.g., by shorting pins 3-4).  When the signal is pulled low during the execution of a program, the controller resets itself back to the beginning of the program.  Program execution can be resumed by either a software start command or by a hardware trigger.

SMA Connector labeled “Ext_Trig”

This is an input connector, and its function depends on the design.   On some designs, a TTL signal applied to this connector will allow the user to modulate the DDS RF output amplitude.   The active pin is pulled high by a 10 kohm on-board resistor.  Setting Ext_Trig to a logical zero will attenuate the RF output amplitude by a preset value, as requested by the user upon ordering.   This feature may be utilized for bi-level decoupling applications in NMR. 
SMA Connector labeled “Ext_Clk”


This SMA connector is used as an external clock input.  Before attaching the external clock, the internal clock must be removed from its socket.  The internal clock’s orientation should be noted (if the internal clock is reconnected, it must be inserted in the same orientation or board damage may occur).  Also, the external clock must be a 3.3 V TTL signal.  Another requirement is that a 50 ohm resistor be soldered directly on the board on R401 or use a T connector with a 50 ohm terminator connected directly to SMA400.  Failure to follow the above requirements could result in damaging the board.    

Appendix I: Sample C program

Example Program

// Example2.cpp

//

// SpinCore Technologies, Inc.

// October 2002

// http://www.spincore.com

//

// The following program code uses C Functions from ‘pbdfuncts’ to

// generate and execute a pulse sequence on the PulseBlasterDDS board.

// Be sure to include the DLL (pbd03pc.dll), the library file

// (pbd03pc.lib), the header files (pbd03pc.h and pbdfuncs.h), and source // file (pbdfuncs.cpp) in the working directory of your C compiler
.

#include <stdio.h>

#include "pbdfuncs.h"

#include "pbdfuncs.cpp"

void main(void)

{


// Located and Initialize Control of the PulseBlasterDDS


if (pb_init() != 0)



printf("--- Error Initializing PulseBlasterDDS ---\n");


// Set clock frequency


set_clock(50);


// Start programming the frequency registers


start_programming(FREQ_REGS);


// Program the registers in order from 0 to 15


// Valid units are MHz, kHz, and Hz - default is MHz


set_freq(1*MHz); // Set register 0


set_freq(2*MHz); // Set register 1


set_freq(3*MHz); // Set register 2


set_freq(4*MHz); // Set register 3


set_freq(5*MHz); // Set register 4 


set_freq(6*MHz); // Set register 5


set_freq(7*MHz); // Set register 6


set_freq(8*MHz); // Set register 7


set_freq(9*MHz); // Set register 8


set_freq(10*MHz); // Set register 9


set_freq(11*MHz); // Set register 10


set_freq(12*MHz); // Set register 11


set_freq(13*MHz); // Set register 12


set_freq(14*MHz); // Set register 13


set_freq(15*MHz); // Set register 14


set_freq(16*MHz); // Set register 15


// Start programming the phase registers


start_programming(PHASE_REGS);


// Program the registers in order from 0 to 15


// Units are in degrees


set_phase(0); // Set register 0


set_phase(22.5); // Set register 1


set_phase(45); // Set register 2


set_phase(67.5); // Set register 3


set_phase(90); // Set register 4


set_phase(112.5); // Set register 5


set_phase(135); // Set register 6


set_phase(157.5); // Set register 7


set_phase(180); // Set register 8


set_phase(202.5); // Set register 9


set_phase(225); // Set register 10


set_phase(247.5); // Set register 11


set_phase(270); // Set register 12


set_phase(292.5); // Set register 13


set_phase(315); // Set register 14


set_phase(337.5); // Set register 15


//Begin pulse program


start_programming(PULSE_PROGRAM);


int start, loop, sub; // define instruction labels


sub = 5; // Since we are going to jump forward in our program, we 




// need to define this variable by hand.  Instructions 




// start at 0 and count up


// Instruction format

// int pb_inst(int freq, int phase, int RF_Output_Enable, int flags,

// int inst, int inst_data, int length)


// Instruction 0 - Jump to Subroutine at Instruction 4 in 1us


start = pb_inst(0,0,ANALOG_OFF,0,0,0x3FF,JSR,sub,1*us);


// Instruction 1 - Beginning of Loop (Loop 3 times).  Continue to 


// next instruction in 1us


loop = pb_inst(0,0,ANALOG_OFF,0,0,0x0,LOOP,3,1*us);


// Instruction 2 - End of Loop.  Return to beginning of loop or 


// continue to next instruction in 1us


pb_inst(0,0,ANALOG_ON,0,0,0x0,END_LOOP,loop,1*us); 


// Instruction 3 - Stay here for (5*1us) then continue to


// next instruction


pb_inst(0,0,ANALOG_OFF,0,0,0x0,LONG_DELAY,5,1*us);


// Instruction 4 - Branch to "start" (Instruction 0) in 1us


pb_inst(0,0,ANALOG_OFF,0,0,0x0,BRANCH,start,1*us);


// Instruction 5 - Continue to next instruction in 2us


pb_inst(0,0,ANALOG_OFF,0,0,0x0,CONTINUE,0,2*us);


// Instruction 6 - Return from Subroutine to Instruction 1 in 2us


pb_inst(0,0,ANALOG_ON,0,0,0x0,RTS,0,2*us);


// End of programming registers and pulse program


stop_programming();


// Start execution of the pulse program


start_pb();


pb_close();

}

Appendix II: Programming the PulseBlasterDDS Using Direct Outputs

If you do not wish to use the provided C functions, you can take advantage of the PBD03PC_outp() function in your own applications.  An explanation of using this function to program the board is included below.

Using DLL Functions to Send Instructions

The provided driver DLL, pbd03pc.dll, provides three functions necessary for programming the PulseBlasterDDS:

int PBD03PC_Init();
This function locates and initializes the PulseBlasterDDS PCI board.  It returns a 0 upon successful completion, or a negative number for an error.

int PBD03PC_outp(unsigned short address, int data);
This function sends 8 bits of data to the PCI board, at a specified address offset.  (Only the lowest 8 bits of the integer value are sent)

int PBD03PC_Close();
This function releases control of the PulseBlasterDDS PCI board.  It returns a 0 upon successful completion, or a negative number for an error.

Building Instructions Using the DLL Functions

· To send instructions to the PulseBlasterDDS, the programmer must first call the PBD03PC_Init() function to locate and initialize the PulseBlasterDDS for sending instructions


· Instructions are then build using the PBD03PC_outp(unsigned short address, int data) function, 8 bits at a time.  For example, 10 PBD03PC_outp function calls are made to construct one 80-bit instruction in accordance with the Instruction Set Architecture


· Finally, the PBD03PC_Close() instruction is called to signal an end to programming, and to release control of the PulseBlasterDDS
Programming Information

Initialization of the PulseBlasterDDS for operation involves a minimum of four steps.  The steps are as follows:

1) Send LOAD NUMBER OF BYTES PER WORD instruction.

2) Send SELECT PERIPHERAL DEVICE instruction.

3) Send CLEAR ADDRESS COUNTER instruction.

4) Loading data to memory.*

5) Send PROGRAMMING FINISHED instruction.

* Once the board has been programmed, this step is optional.  If you are just restarting the pulse program after a STOP command or an HARDWARE_RESET has stopped the program, you can skip this step during re-initialization.

If these five commands are not sent from a PC, the PulseBlasterDDS will not run as desired.  All five instructions are required as an attempt to ensure that the device has been programmed before it can be armed.  The first time the system is used, the loading of the memory with data has to be performed.  Upon reset, four instructions must be executed to restart the device again.  

WE for Peripherals: This register is used to select the peripheral that is to be programmed.  The value of this register that is used to select program memory is always zero and this is the default value for the register.  A complete listing of the values and the associated hardware that can be programmed when appropriately set is listed below in Table A1.1.  By selecting 0xFF for this value, no device is selected and the initial value of the output flags may be set.  (*Note that the clock signal to the FF must still be transitioned, which may be accomplished by writing to the base output port + 5 twice in succession.)




[image: image2.wmf]WE Register

Value (hex)

Program Memory

0

Program Frequency Registers

1

Program Phase Registers

2

Change Flags Only

FF


Table A1.1: Peripheral List

CLEAR ADDRESS COUNTER: The Address Counter is used to manufacture the memory address.  The Address Counter is not loadable; it can only be cleared and started at zero.  It is not possible to load a particular section of memory.  All loads must start from either the beginning of memory, or wherever the Address Counter left off.

Flag Initialization Strobe: The output flags of the PulseBlasterDDS can be programmed while the device is in a reset state.  This is useful to initialize flags after powering-up and to reset flags to a known state if a program must be aborted.  Writing to the Flag Initialization Strobe register will toggle the line used to clock data into the output latches.

LOAD_MEMORY: This instruction is used to specify data that should be used to program the memory used by the device.  Since the incoming data is taken only one byte at a time, the IBC must reconstruct the data word to be programmed.  The data word is reconstructed in the IBC most significant byte first.

PROGRAMMING FINISHED: This instruction enables the pattern generator of the PulseBlasterDDS.  This instruction prevents the pattern generator from accepting a hardware trigger or software start command before the device has been programmed.  Once the design has been programmed, the PROGRAMMING FINISIHED command must be sent to arm the device for operation.  After the pattern generator has been armed, any hardware trigger or software start command will cause the system to start operation. The PulseBlasterDDS can be reset by issuing the DEVICE_RESET command.  This will internally clear the PROGRAMMING FINISHED instruction and prevent the pattern generator from operating again until the IBC has been re-initialized.

In order to select each of the commands mentioned about, you write to the port base + offset.  A table of each offset’s meaning is included below.

	Offset
	Command
	Function

	0
	DEVICE_RESET
	Stops Pulse Program

	1
	DEVICE_START
	Starts Pulse Program (only when in initialized state)

	2
	SELECT_BPW
	Selects number of BPW (10 for instructions)

	3
	SELECT_DEVICE
	Using Table 2, selects internal device to be programmed

	4
	CLEAR_ADDRESS_COUNTER
	Resets internal memory address counter

	5
	FLAG_STROBE
	Strobes output clock signal to preset digital outputs

	6
	DATA_TRANSFER
	Data to be written to internal device

	7
	PROGRAMMING_FINISHED
	Sets device in initialized state


Table A1.2: Port Offset Command Functions
Example Program

The following is an example of the output sequence to program the PulseBlasterDDS.  You must use the PBD03PC_outp(addr, data) function from the provided DLL in order to use this method.

//
Initialization:

PBD03PC_outp(0,0);


//(Issue device reset)

PBD03PC_outp(2,4);


//(Select number of bytes per word)

PBD03PC_outp(3,0xFF);


//(Select device to program (Flag initial values))

PBD03PC_outp(4,0);


//(Reset address counter)

//
Set initial flag values

//
Values for this example are "0x000000f0"

PBD03PC_outp(6,0);


//(Data transfer)

PBD03PC_outp(6,0);


//(Data transfer)

PBD03PC_outp(6,0);


//(Data transfer)

PBD03PC_outp(6,0xF0);


//(Data transfer)

PBD03PC_outp(5,0);


//(Clock data into external buffer)

PBD03PC_outp(5,0);


//(Return clock signal to low)

//
Set up DDS frequency registers

PBD03PC_outp(0,0);


//(Issue device reset)

PBD03PC_outp(2,4);


//(Select number of bytes per word)

PBD03PC_outp(3,1);


//(Select device to program (DDS Frequency Registers))

PBD03PC_outp(4,0);


//(Reset address counter)

/*


DDS Register Values


Reg0 = 051EB852 (1 MHz)


Reg1 = 0A3D70A4 (2 MHz)


.


.


Reg15 = 147AE148 (4 MHz)


Formula for finding these values:


REG0 = DESIRED_FREQUENCY * 232 / PBDDS_CLOCK



= 1 MHz * 232 / 50 MHz = 858993459.2 = 0x051EB852

*/

PBD03PC_outp(6,0x05);

//(Data Transfer - Byte 3 of Reg0)

PBD03PC_outp(6,0x1E);

//(Data Transfer - Byte 2 of Reg0)

PBD03PC_outp(6,0xB8);

//(Data Transfer - Byte 1 of Reg0)

PBD03PC_outp(6,0x52);

//(Data Transfer - Byte 0 of Reg0)

PBD03PC_outp(6,0x05);

//(Data Transfer - Byte 3 of Reg1)

PBD03PC_outp(6,0x3D);

//(Data Transfer – Byte 2 of Reg1)

PBD03PC_outp(6,0x70);

//(Data Transfer - Byte 1 of Reg1)

PBD03PC_outp(6,0xA4);

//(Data Transfer - Byte 0 of Reg1)

.

.

PBD03PC_outp(6,0x14);

//(Data Transfer - Byte 3 of Reg15)

PBD03PC_outp(6,0x7A);

//(Data Transfer - Byte 2 of Reg15)

PBD03PC_outp(6,0xE1);

//(Data Transfer - Byte 1 of Reg15)

PBD03PC_outp(6,0x48);

//(Data Transfer - Byte 0 of Reg15)

//
Set up phase registers

PBD03PC_outp(0x0);


//(Issue device reset)

PBD03PC_outp(2,4);


//(Select number of bytes per word)

PBD03PC_outp(3,2);


//(Select device to program (TX Phase Registers))

PBD03PC_outp(4,0);


//(Reset address counter)

/*


Phase Register Values


Reg0 = 093E93E9 (13 degrees)


Reg1 = 40000000 (90 degrees)



.


.


Reg15 = E0000000(315 degrees)


Formula for finding these values:


REG0 = DESIRED_PHASE * 232 / 360



= 13 * 232 / 360 = 155096041.24 ( 0x093E93E9


* Note that only the top 3 hex digits (0x093) are actually used in programming the phase register.  The other digits must still be sent so that the DDS word remains 32 bits long.

*/

PBD03PC_outp(6,0x09);

//(Data Transfer - Byte 3 of Reg0)

PBD03PC_outp(6,0x3E);

//(Data Transfer - Byte 2 of Reg0)

PBD03PC_outp(6,0x93);

//(Data Transfer - Byte 1 of Reg0)

PBD03PC_outp(6,0xE9);

//(Data Transfer - Byte 0 of Reg0)

PBD03PC_outp(6,0x40);

//(Data Transfer - Byte 3 of Reg1)

PBD03PC_outp(6,0x00);

//(Data Transfer - Byte 2 of Reg1)

PBD03PC_outp(6,0x00);

//(Data Transfer - Byte 1 of Reg1)

PBD03PC_outp(6,0x00);

//(Data Transfer - Byte 0 of Reg1)

.

.

PBD03PC_outp(6,0xE0);

//(Data Transfer - Byte 3 of Reg15)

PBD03PC_outp(6,0x00);

//(Data Transfer - Byte 2 of Reg15)

PBD03PC_outp(6,0x00);

//(Data Transfer - Byte 1 of Reg15)

PBD03PC_outp(6,0x00);

//(Data Transfer - Byte 0 of Reg15)

//
Pulse Program Setup

PBD03PC_outp(0,0);


//(Issue Device Reset)

PBD03PC_outp(2,10);

//(Select number of bytes per word)

PBD03PC_outp(3,0);


//(Select device to program (RAM))

PBD03PC_outp(4,0);


//(Reset address counter)

//Send pulse program

PBD03PC_outp(6,0x18);

//(Byte 9 of first instruction)

PBD03PC_outp(6,0xFF);

//(Byte 8 of first instruction)

PBD03PC_outp(6,0xFF);

//(Byte 7 of first instruction)

PBD03PC_outp(6,0x00);

//(Byte 6 of first instruction)

PBD03PC_outp(6,0x00);

//(Byte 5 of first instruction)

PBD03PC_outp(6,0x00);

//(Byte 4 of first instruction)

PBD03PC_outp(6,0x00);

//(Byte 3 of first instruction)

PBD03PC_outp(6,0x00);

//(Byte 2 of first instruction)

PBD03PC_outp(6,0x00);

//(Byte 1 of first instruction)

PBD03PC_outp(6,0x07);

//(Byte 0 of first instruction)

PBD03PC_outp(6,0xXX);

//(Byte 9 of second instruction)

/*

Continue this process for all instructions.  The

explanation of how to create the 80 bit instruction

words is included below.  When finished with all

instructions, continue with the sequence below.

*/

PBD03PC_outp(7,0);


//(Programming Finished)

/*

Only execute the following command when you are ready

for the program to start running.

*/

PBD03PC_outp(1,0);


//(Start pulse program)
Contact Information

Phone



(352) 271-7383


FAX




(352) 371-8679


Email




sales@spincore.com


Web




http://www.spincore.com/

Product URL:

http://www.pulseblaster.com/CD/PulseBlasterDDS/PCI/PBD03PC



� EMBED Word.Picture.8  ���


























� These functions and library files have been generated and tested with the MS Visual Studio 6 environment.  Support and updated functions/DLLs for other environments may be provided upon request if available.


� These functions and library files have been generated and tested with the MS Visual Studio 6 environment.  Support and updated functions/DLLs for other environments may be provided upon request if available.





21

4/4/2005

_1092836160.xls
Sheet1

		WE Register		Value (hex)

		Program Memory		0

		Program Frequency Registers		1

		Program Phase Registers		2

		Change Flags Only		FF






_1093351654.doc
[image: image1.png]DDS Core

Tomerically
Controlled
Oseillator

Reference
Oscillator

Outputs

Output and Control Register

2002 SpinCore Technologies, Inc.

User Control :
http://www spincore.com








