PulseBlasterDDS™

Model DDS-III
(PCI Board SP3)

Owner’s Manual

SpinCore Technologies, Inc.

http://www.spincore.com

Congratulations and thank you for choosing a design from SpinCore Technologies, Inc.

We appreciate your business!

At SpinCore we try to fully support the needs of our customers.  If you are in need of assistance, please contact us and we will strive to provide the necessary support.

© 2000-2004 SpinCore Technologies, Inc. All rights reserved.
SpinCore Technologies, Inc. reserves the right to make changes to the product(s) or information herein without notice. PulseBlasterDDS™, PulseBlaster™, SpinCore, and the SpinCore Technologies, Inc. logos are trademarks of SpinCore Technologies, Inc. All other trademarks are the property of their respective owners.

SpinCore Technologies, Inc. makes every effort to verify the correct operation of the equipment.  This equipment version is not intended for use in a system in which the failure of a SpinCore device will threaten the safety of equipment or person(s).

Table of Contents

I. Introduction	3

Product Overview	3

Board Architecture	4

Block Diagram	4

Output signals	4

Timing characteristics	5

Instruction set	5

External triggering	5

Status Readback	5

Summary	5

Specifications	6

DDS Specifications	6

TTL Specifications	6

Common Parameters (DDS and TTL Specifications)	6

Pulse Program Control Flow (Common)	6

II. Installation	7

Installing the PulseBlasterDDS Driver	7

For Windows XP	7

Initializing Control of the PulseBlasterDDS	8

III. Programming the PulseBlasterDDS	9

Instruction Set Architecture	9

Machine-Word Definition	9

Breakdown of 80-bit Instruction Word	9

Using C Functions to Program the PulseBlasterDDS	11

Example Use of C Functions	13

IV. Connecting to the PulseBlasterDDS Board	15

Connector Information	15

SMA Connectors labeled DAC_OUT_0, DAC_OUT_1, and DAC_OUT_2	15

DB-25 - TTL Output Signal Bits	15

IDC Connector Status - Pin Assignments	16

Header JP100	16

SMA Connector labeled “SMA0”	16

SMA Connector labeled “SMA400”	17

Appendix I: Sample C program	17

Example Program	17

Appendix II: Programming the PulseBlasterDDS Using Direct Outputs	20

Using DLL Functions to Send Instructions	20

Building Instructions Using the DLL Functions	20

Programming Information	20

Example Program	22

Contact Information	25

# I. Introduction

## Product Overview

The PulseBlasterDDS series of Intelligent Pattern and Waveform Generation boards from SpinCore Technologies, Inc., couples SpinCore’s unique Intelligent Pattern Generation processor core, dubbed PulseBlaster, with Direct Digital Synthesis (DDS) for use in system control and pulse generation.

The PulseBlaster’s state-of-the-art timing processor core provides all the necessary timing control signals required for overall system control and pulse synchronization.   By adding DDS features, PulseBlasterDDS can now provide not only digital (TTL) but also analog output signals, meeting high-performance and high-precision complex excitation/stimuli needs of demanding users.

PulseBlasterDDS provides users the ability to control their systems through the generation of fully synchronized (digital and analog) excitation pulses from a small form factor PC board, providing users a compelling price/performance proposition unmatched by any other device on the market today.  Figure 1 presents sample capabilities of the board.

Figure 1: Sample PulseBlasterDDS output capabilities

## Board Architecture

### Block Diagram

Figure 2 presents the general architecture of the PulseBlasterDDS system.  The two major building blocks are the DDS Core and the Pulse Programming and Timing Processor Core (PP Core). The DDS Core contains a numerically controlled oscillator and has 16 programmable frequency registers that are under the pulse program control.  Prior to gating, the DDS signal can be phase offset by one of two sets of 16 programmable phase registers.  The PP Core controls the timing of the gating pulses and provides the necessary control signals for frequency and phase registers.  The DDS and PP cores have been integrated onto a single silicon chip.  High performance DAC chips and high current output amplifiers complement the design.  User control to the system is provided through the host-programming interface over the PCI bus.

Figure 2: PulseBlasterDDS board architecture

### Output signals

The PulseBlasterDDS comes with three analog output channels configured to output radio-frequency (RF/IF) pulses, and 10 digital output signal lines (one of the output lines has a dual use and

functions as a phase reset for the DDS generator). The frequency and phase of the RF pulses generated by the DDS are under the control of the user and are specified through software programming.  The phase of the numerically controlled oscillator can be reset on demand within the pulse program.  PulseBlasterDDS provides the ability to gate the output of the DDS channels allowing for independent pulsed RF operation.  With digital sampling rate of 100 MHz (max. reference clock frequency), the maximum theoretical output frequency is 50 MHz (the Nyquist Theorem).   The analog output signal is available on an on-board SMA connector.  The output impedance of the analog signal is 50-ohms.  There are no interpolating filters on board.

The 10 individually controlled digital (TTL/CMOS) output bits are capable of delivering 25 mA per bit and have an output voltage of 3.3V. These signals are available on the PC bracket-mounted DB-25 connector.  Setting output bit 10 high via the output control word also resets the phase of the RF waveforms for phase coherent switching, and can be used to generate a constant voltage on the DACs.

### Timing characteristics

PulseBlasterDDS’s timing controller can accept either an internal (on-board) crystal oscillator or an external frequency source of up to 100 MHz. The innovative architecture of the timing controller allows the processing of either simple timing instructions (delays of up to 232 = 4,294,967,296 clock cycles), or double-length timing instructions (up to 252 clock cycles long – nearly 2 years with a 100 MHz clock!).  Regardless of the type of timing instruction, the timing resolution remains constant for any delay – just one clock period (e.g., 10 ns for a 100 MHz clock).

The timing controller has a very short minimum delay cycle – only nine clock periods. This translates to a 90 ns minimum pulse/delay/update with a 100 MHz clock.

### Phase Coherent Switching

The board allows for phase continuous and/or phase coherent switching.  In addition, the DDS can be reset to zero whenever a new RF pulse is started.  Consult the explanation of the flags parameter to the pb_inst instruction on page 12 for implementing the phase reset.

### Instruction set

PulseBlasterDDS’ design features a set of commands for highly flexible program flow control. The micro-programmed controller allows for programs to include branches, subroutines, and loops at up to 8 nested levels – all this to assist the user in creating dense pulse programs that cycle through repetitious events, especially useful in numerous multidimensional spectroscopy and imaging applications.

### External triggering

PulseBlasterDDS can be triggered and/or reset externally via dedicated hardware lines.  The two separate lines combine the convenience of triggering (e.g., in cardiac gating) with the safety of the "stop/reset" line.  The required control signals are “active low” (or short to ground).

### Status Readback

The status of the program can be read in hardware or software.  The hardware status output signals consist of five IDC connector pins labeled “Status”.  The same output can be read through software using C.  See section IV (Connecting to the PulseBlaster Board, page 16) for more detail about the hardware lines and section III (Programming the PulseBlaster, page 11) for more detail about the C function status_readback().

### Summary

PulseBlasterDDS is a versatile, high-performance pulse/pattern TTL and RF/IF generator operating at speeds of up to 100 MHz and capable of generating pulses/delays/intervals ranging from 90 ns to over 2 years per instruction. It can accommodate pulse programs with highly flexible control commands of up to 32k program words. Its high-current output logic bits are independently controlled with a voltage of 3.3 V.  The output impedance of the analog channel is 50-ohms.

## Specifications

### DDS Specifications

- 100 MHz reference clock oscillator (other frequencies available upon request)

- 0.047 Hz frequency resolution (32 bits)

- 16 loadable frequency registers for agile frequency modulation/switching/selection (32 bits each)

- Two sets of 16 loadable phase-offset registers for agile phase modulation/switching/selection (12 bits each)

- 0.09 phase resolution (12 bits)

- 40 ns phase switching latency

- 40 ns frequency switching latency (phase continuous)

- phase coherent switching

- 10 dBm RF output power

- 50 ohm output impedance

- SMA connectors

- 30 MHz 3dB bandwidth

- RF Output capable of outputting DC at programmed output level (using phase offset)

### TTL Specifications

- 10 individually controlled digital output lines (TTL levels; one of the output lines has a dual use and functions as a phase reset for the DDS generator)

- variable pulses/delays for every TTL line

- 25 mA output current per TTL line

- output lines can be combined to increase the max. output current

### Common Parameters (DDS and TTL Specifications)

- 90 ns shortest pulse/interval per instruction

- 2 years longest pulse/interval per instruction

- 10 ns pulse/interval resolution

- RF and TTL pulses are synchronized

- 32k max. memory space

- external  triggering and reset – TTL levels

### Pulse Program Control Flow (Common)

- loops, nested 8 levels deep

- 20 bit loop counters (max. 1,048,576 repetitions)

- subroutines, nested 8 levels deep

- wait for trigger - 80 ns latency, adjustable to 2 years in duration

- Approximately 2 MHz max. re-triggering frequency (based on the latency of the WAIT opcode)

# II. Installation

## Installing the PulseBlasterDDS Driver

- Go to http://www.pulseblaster.com/CD/PulseBlasterDDS/PCI/SP3 and download sp3.zip.

- Unzip the files to their own directory.

- Turn off your computer.

- Insert the PulseBlasterDDS board into an empty PCI slot.  Secure the bracket firmly with a screw.

- Turn on your computer.

## For Windows XP

6.  After booting, the “Found New Hardware Wizard” should appear.  Choose “Install from a list or specific location” and click Next

7.  Choose “Include this location in the search” and browse to the directory you unzipped the drivers to.  Click next.

8.  While windows installs the driver, a “Files Needed” dialog may pop up.  Choose the directory you unzipped the drivers to, and click ok.

9.  When finished, you should see this window.

NOTE: On some systems after you install your PulseBlasterDDS board you may need to run “Install.bat” located in the “post_installation_files.zip” file located at http://www.pulseblaster.com/CD/PulseBlasterDDS/PCI/SP3/old_version/post_installation_files.zip

in order for your board to work.

You are now ready to control the PulseBlasterDDS board

## Initializing Control of the PulseBlasterDDS

3. Run the included “SP3_Test.exe”.

If equipped with a 100 MHz reference clock oscillator, the board should now output a 6.250MHz sine wave on the SMA connectors labeled “DAC_OUT_0”, “DAC_OUT_1”, and “DAC_OUT_2”.  The 10 TTL output lines should toggle every second.

The PulseBlasterDDS board is now ready for use!

# III. Programming the PulseBlasterDDS

## Instruction Set Architecture

### Machine-Word Definition

The PulseBlaster pulse timing and control processor implements an 80-bit wide Very Long Instruction Word (VLIW) architecture. The VLIW memory words have specific bits/fields dedicated to specific purposes, and every word should be viewed as a single instruction of the micro-controller. The maximum number of instructions that can be loaded to on-board memory is 32k. The execution time of instructions can be varied and is under (self) control by one of the fields of the instruction word – the shortest being five clock cycles (for 512 memory-word models) and the longest being 2^52 clock cycles.  All instructions have the same format and bit length, and all bit fields have to be filled.  Figure 3 shows the fields and bit definitions of the 80-bit instruction word.

Bit Definitions for the 80-bit Instruction Word (VLIW)

Output/Control Word   |    Data Field    |     OP Code    |    Delay Count

(24 bits)             (20 bits)          (4 bits)         (32 bits)

Figure 3: Bit definitions of the 80-bit instruction/memory word

### Breakdown of 80-bit Instruction Word

The 80-bit VLIW is broken up into 4 sections

1. Output Pattern and Control Word - 24 bits

2. Data Field - 20 bits

3. OP Code - 4 bits

4. Delay Count - 32 bits

#### Output Pattern and Control Word

Please refer to Table 1, next page, for output pattern and control bit assignments of the 24-bit output/control word.

| Bit # | Function | Bit # | Function |
| --- | --- | --- | --- |
| 23 | Selects Frequency Register (bit 3) | 11 | Selects Phase Register for SMA connectors labeled DAC_OUT_2 and DAC_OUT_0 (bit 0) |
| 22 | Selects Frequency Register (bit 2) | 10 | Output Enable for SMA connectors labeled DAC_OUT_1 and DAC_OUT_0 (0 = on, 1 = off) |
| 21 | Selects Frequency Register (bit 1) | 9 | RF phase reset for phase coherent switching, also routed to Output Connector DB25 pin 19 |
| 20 | Selects Frequency Register (bit 0) | 8 | Output Connector DB25 pin 7 |
| 19 | Selects Phase Register for SMA connector labeled DAC_OUT_1 (bit 3) | 7 | Output Connector DB25 pin 8 |
| 18 | Selects Phase Register for SMA connector labeled DAC_OUT_1 (bit 2) | 6 | Output Connector DB25 pin 21 |
| 17 | Selects Phase Register for SMA connector labeled DAC_OUT_1 (bit 1) | 5 | Output Connector DB25 pin 22 |
| 16 | Selects Phase Register for SMA connector labeled DAC_OUT_1 (bit 0) | 4 | Output Connector DB25 pin 10 |
| 15 | Output Enable for SMA connector labeled DAC_OUT_2 (0 = on, 1 = off) | 3 | Output Connector DB25 pin 11 |
| 14 | Selects Phase Register for SMA connectors labeled DAC_OUT_2 and DAC_OUT_0 (bit 3) | 2 | Output Connector DB25 pin 24 |
| 13 | Selects Phase Register for SMA connectors labeled DAC_OUT_2 and DAC_OUT_0 (bit 2) | 1 | Output Connector DB25 pin 25 |
| 12 | Selects Phase Register for SMA connectors labeled DAC_OUT_2 and DAC_OUT_0 (bit 1) | 0 | Output Connector DB25 pin 13 |

Table 1: Output Pattern and Control Word Bits

#### Data Field and Op Code

Please refer to Table 2 for information on the available operational codes (OpCode) and the associated data field functions (the data field's function is dependent on the Op Code)

| Op Code | Instruction | Data Field | Function |
| --- | --- | --- | --- |
| 0 | CONTINUE | Ignored | Program execution continues to next instruction |
| 1 | STOP | Ignored | Stop execution of program (*Note all TTL values remain from previous instruction, and analog outputs turn off) |
| 2 | LOOP | Number of desired loops. This value must be greater than or equal to 1. | Specify beginning of a loop. Execution continues to next instruction. Data used to specify number of loops |
| 3 | END_LOOP | Address of beginning of loop | Specify end of a loop. Execution returns to begging of loop and decrements loop counter. |
| 4 | JSR | Address of first subroutine instruction | Program execution jumps to beginning of a subroutine |
| 5 | RTS | Ignored | Program execution returns to instruction after JSR was called |
| 6 | BRANCH | Address of next instruction | Program execution continues at specified instruction |
| 7 | LONG_DELAY | Number of desired loops. This value must be greater than or equal to 2. | For long interval instructions. Data field specifies a multiplier of the delay field. Execution continues to next instruction |
| 8 | WAIT | Ignored | Program execution stops and waits for software or hardware trigger. Execution continues to next instruction after receipt of trigger. The latency is equal to the delay value entered in the WAIT instruction line plus a fixed delay of 6 clock cycles. |

Table 2: Op Code and Data Field Description

#### Delay Count

The value of the Delay Count field (a 32-bit value) determines how long the current instruction should be executed.  The allowed minimum value of this field is 0x6 for the 32k memory models.  The timing controller has a fixed delay of three clock cycles and the value that one enters into the Delay Count field should account for this inherent delay.

## Using C Functions to Program the PulseBlasterDDS

A series of functions have been written to control the board and facilitate the construction of pulse program instructions.  The functions also allow the programmer to set the DDS frequency and phase registers.

In order to use these functions, the DLL (pbd03pc.dll), the library file (pbd03pc.lib), the header files (pbd03pc.h and pbdfuncs.h), and source file (pbdfuncs.cpp) must be in the working directory of your C compiler.

int pb_init();

Initializes PulseBlasterDDS board.  Needs to be called before calling any functions using the PulseBlasterDDS.  Returns a negative number on an error or 0 on success.

int pb_close();

Releases PulseBlasterDDS board.  Needs to be called as last command in pulse program.  Returns a negative number on an error or 0 on success.

void set_clock(double clock_freq);

Used to set the clock frequency of the board.  The variable clock_frequency is specified in MHz when no units are entered.  Valid units are MHz, kHz, and Hz.  The default clock value is 50MHz.  You only need to call this function if you are not using a 50MHz board.

int start_programming(int device);

Used to initialize the system to receive programming information.  It accepts a parameter referencing the target for the instructions.  Valid values for device are PULSE_PROGRAM, FREQ_REGS, PHASE_REGS_0, and PHASE_REGS_1.  PHASE_REGS_0 programs the phase registers for the DDS output on SMA connectors labeled DAC_OUT_0 and DAC_OUT_2.  PHASE_REGS_1 programs the phase registers for the DDS output on SMA connector labeled DAC_OUT_1.  The function returns a 0 on success or a negative number on an error.

int set_freq(double freq);

Used to set the values in the frequency registers.  Should only be called after start_programming(FREQ_REGS) has been called.  Registers are programmed one at a time, starting at 0 and incrementing each time this function is called.  It accepts the value for the frequency register with a default unit of MHz.  Valid units are MHz, kHz, Hz.  It returns a 0 on success or a negative number on an error.

int set_phase(double phase);

Used to set the values in the phase registers.  Should only be called after start_programming(PHASE_REGS_0) or start_programming(PHASE_REGS_1) has been called.  Registers are programmed one at a time, starting at 0 and incrementing each time this function is called.  It accepts the value for the phase register in degrees.  It returns a 0 on success or a negative number on an error.

int pb_inst(int freq, int phase_SMA_1, int tx_output_enable,

int phase_SMA_0, int rx_output_enable, int flags, int inst,

int inst_data, double length);

Used to send one instruction of the pulse program.  Should only be called after start_programming(PULSE_PROGRAM) has been called.  It returns a negative number on an error, or the instruction number upon success.  If the function returns –99, an invalid parameter was passed to the function.  Instructions are numbered starting at 0.

int freq – Selects the frequency register to be used.  Valid range is from 0 to 16

int phase_SMA_1 – Selects the phase register to be used from the phase registers programmed using start_programming(PHASE_REG_1).  This is the DDS output on SMA connector labeled DAC_OUT_1.  Valid range is from 0 to 16

int tx_output_enable – Determines whether analog output is generating a sinusoid or is at ground for TX output on SMA connector labeled DAC_OUT_2.  Valid values are TX_ANALOG_ON and TX_ANALOG_OFF

int phase_SMA_0 – Selects the phase register to be used from the phase registers programmed using start_programming(PHASE_REG_0).  This is the DDS output on SMA connectors labeled DAC_OUT_0 and DAC_OUT_2.  Valid range is from 0 to 16

int rx_output_enable – Determines whether analog output is generating a sinusoid or is at ground for RX output on SMA connectors labeled DAC_OUT_0 and DAC_OUT_1.  Valid values are RX_ANALOG_ON and RX_ANALOG_OFF

int flags – determines state of each TTL output bit.  Valid values are 0x0 to 0x3FF.  For example, 0x010 would correspond to bit 5 being on and all other bits being off.  Bit 10, corresponding to hexadecimal value 0x200, is used to reset the phase of the numerically controlled oscillator.  This results in phase coherent switching.  The numerical oscillator will pause for the length of the instruction, resulting in a constant voltage at the DAC.  The voltage level depends on the value of the phase register in use.

int inst – determines which type of instruction is to be executed. Please see Table 2 for details.

int inst_data – data to be used with the previous inst field.  Please see Table 2 for details.

double length – duration of this pulse program instruction, specified in ns.

This function has been overloaded to accommodate TTL-only programs.  When using the shorter version of the function, the RF channel has its output set to ground. The overloaded form follow:

TTL Only:

int pb_inst(int flags, int inst, int inst_data, double length);

int stop_programming();

Used to tell that programming the board is complete.  Board execution cannot start until this command is received.  It returns a 0 on success or a negative number on an error.

int start_pb();

Once board has been programmed, this instruction will start execution of pulse program.  It returns a 0 on success or a negative number on an error.

int stop_pb();

Stops output of board.  Analog output will return to ground, and TTL outputs will remain in the state they were in when stop command was received.  It returns a 0 on success or a negative number on an error.

int status_readback();

Reads the status of the board and returns an integer whose bit representation corresponds to the status signals described in section IV.  Bit zero is stopped; bit one is reset; bit two is running; bit three is waiting.

### Example Use of C Functions

// Example1.cpp

//

// SpinCore Technologies, Inc.

// May 2004

// http://www.spincore.com

//

// The following program code uses C Functions from 'pbdfuncs' to

// generate and execute a pulse sequence on the PulseBlasterDDS board.

// Be sure to include the DLL (pbd03pc.dll), the library file

// (pbd03pc.lib), the header files (pbd03pc.h and pbdfuncs.h), and source

// file (pbdfuncs.cpp) in the working directory of your C compiler .

#include "pbdfuncs.h"

#include "pbdfuncs.cpp"

#include "PBD03PC.h"

#include <stdio.h>

void main(void)

{

int start;

// Locates & Initializes the PulseBlasterDDS Board

pb_init();

// Set Because the Board Operates at 100MHz

set_clock(100);

// Prepare the Board to Receive Freqeuncy Values

start_programming(FREQ_REGS);

// Load Frequency Register 0

set_freq(1.054);

// Load Frequency Register 1

set_freq(2);

// Prepare the Board to Receive TX Phase Values

start_programming(PHASE_REGS_1);

// Load Phase Register 0

set_phase(0);

// Load Phase Register 1

set_phase(90);

// Prepare the Board to Receive RX Phase Values

start_programming(PHASE_REGS_0);

// Load Phase Register 0

set_phase(0);

// Load Phase Register 1

set_phase(180);

// Prepare the Board to Receive pulse program instructions

start_programming(PULSE_PROGRAM);

//Instruction 0 - Continue to instruction 1 in 2us

//Freq Reg 0, Phase Reg 1 for DAC_OUT_1, DDS TX Output ON, Phase Reg 0 for

//DAC_OUT_0 and DAC_OUT_2, DDS RX Output ON, Flags = 0x3FF, OPCODE = CONTINUE

start = pb_inst(0, 1, TX_ANALOG_ON, 0, RX_ANALOG_ON, 0x3FF, CONTINUE, 0, 2*us);

// Instruction 1 - Continue to instruction 2 in 4us

//Freq Reg 0, Phase Reg 1 for DAC_OUT_1, DDS TX Output OFF, Phase Reg 0 for

//DAC_OUT_0 and DAC_OUT_2, DDS RX Output ON, Flags = 0x000, OPCODE = CONTINUE

pb_inst(0, 1, TX_ANALOG_OFF, 0, RX_ANALOG_ON, 0x000, CONTINUE, 0, 4*us);

// Instruction 2 - Branch to "start" (Instruction 0) in 2us

//Freq Reg 0, Phase Reg 1 for DAC_OUT_1, DDS TX Output ON, Phase Reg 0 for

//DAC_OUT_0 and DAC_OUT_2, DDS RX Output ON, Flags = 0x000, OPCODE = BRANCH

pb_inst(0, 1, TX_ANALOG_ON, 0, RX_ANALOG_ON, 0x000, BRANCH, start, 2*us);

// Finished Sending Instructions

stop_programming();

// Run the Program

start_pb();

// Release Control of the PulseBlasterDDS Board

pb_close();

}

A more complex program using C Functions is provided in Appendix I.

# IV. Connecting to the PulseBlasterDDS Board

## Connector Information

### SMA Connectors labeled DAC_OUT_0, DAC_OUT_1, and DAC_OUT_2

Outputs DDS signals generated by the user’s Program.  The output impedance is 50 ohms.  Output power is approximately 10 dBm.  Figure 4 is a portion of figure 2 that illustrates which RF output corresponds to which SMA connector.

DAC_OUT_2

DAC

DAC

DAC

RF

(SMA)

Phase0-15

Gate

Gate

Phase0-15

Gate

Tx

Rx

Outputs

DAC_OUT_0

DAC_OUT_1

Figure 4: SMA Connectors

### DB-25 - TTL Output Signal Bits

Outputs TTL signals generated by the user’s Program.  Please consult the table below for bit assignments.

| Pin Assignments |  |  |  |
| --- | --- | --- | --- |
| Pin# | Bit# | Pin# | Bit# |
| 1 | GND | 14 | GND |
| 2 | Reserved | 15 | Reserved |
| 3 | GND | 16 | Reserved |
| 4 | Reserved | 17 | GND |
| 5 | Reserved | 18 | Reserved |
| 6 | GND | 19 | 9 |
| 7 | 8 | 20 | GND |
| 8 | 7 | 21 | 6 |
| 9 | GND | 22 | 5 |
| 10 | 4 | 23 | GND |
| 11 | 3 | 24 | 2 |
| 12 | GND | 25 | 1 |
| 13 | 0 |  |  |

Table 3: Output bits and signals of the PulseBlasterDDS board

### IDC Connector Status - Pin Assignments

The IDC connector labeled Status outputs TTL signals based on status of the user’s program.  Please consult the table below for pin assignments.

| Pin Assignments |  |  |  |
| --- | --- | --- | --- |
| Pin# |  | Pin# |  |
| 1 | Stopped | 5 | Running |
| 2 | GND | 6 | GND |
| 3 | Reset | 7 | Waiting |
| 4 | GND | 8 | GND |

Table 4: Status signals of the PBDDS-III

The status pins correspond to the current state of the pulse program and are defined as follows:

Stopped – Driven high when the PulseBlaster device has encountered a STOP Op Code during program execution and has entered a stopped state.

Reset – Driven high when the PulseBlaster device is in a RESET state and must be reprogrammed before code execution can begin again.

Running – Driven high when the PulseBlaster device is executing a program.  It is low when the PulseBlaster enters either a reset or idle state.

Waiting – the PulseBlaster device has encountered a WAIT Op Code and is waiting for the next trigger (either hardware or software) to resume operation.

### Header JP100

This is an input connector, for hardware triggering (HW_Trigger) and resetting (HW_Reset).

HW_Trigger is pulled high by default, and pin 1 is active (pin 2 = GND).  When a low state is detected (e.g., when shorting pins 1-2), it initiates code execution.   This trigger will also restart execution of a program from the beginning of the code if it is detected after the design has reached an idle state.  The idle state could have been created either by reaching the STOP Op Code of a program, or by the detection of the HW_Reset signal.  When the WAIT Op Code is used in the pulse program, the HW_Trigger will cause the program to continue to the next instruction.

HW_Reset is pulled high by default, and pin 3 is active (pin 4 = GND).  It can be used to halt the execution of a program by pulling it low (e.g., by shorting pins 3-4).  When the signal is pulled low during the execution of a program, the controller resets itself back to the beginning of the program.  Program execution can be resumed by either a software start command or by a hardware trigger.

### SMA Connector labeled “SMA0”

This SMA connector outputs the reference clock as a 3.3 V TTL signal, i.e., it generates positive-only voltage.  The output resembles a square wave if properly terminated.  This signal can be measured with an oscilloscope using either a high impedance probe at the SMA connector or a 50 ohm coaxial line that is terminated.

### SMA Connector labeled “SMA400”

This SMA connector is used as an external clock input.  Before attaching the external clock, the internal clock must be removed from its socket.  The internal clock’s orientation should be noted (if the internal clock is reconnected, it must be inserted in the same orientation or board damage may occur).  Also, the external clock must be a 3.3 V TTL signal.  Another requirement is that a 50 ohm resistor be soldered directly on the board on R401 pads or use a T connector with a 50 ohm terminator connected directly to SMA400.  Failure to follow the above requirements could result in damaging the board.

# Appendix I: Sample C program

## Example Program

// Example2.cpp

//

// SpinCore Technologies, Inc.

// May 2004

// http://www.spincore.com

//

// The following program code uses C Functions from 'pbdfuncts' to

// generate and execute a pulse sequence on the PulseBlasterDDS board.

// Be sure to include the DLL (pbd03pc.dll), the library file

// (pbd03pc.lib), the header files (pbd03pc.h and pbdfuncs.h), and source

// file (pbdfuncs.cpp) in the working directory of your C compiler .

#include "pbdfuncs.h"

#include "pbdfuncs.cpp"

#include "PBD03PC.h"

#include <stdio.h>

void main(void)

{

if (pb_init() != 0)

printf("--- Error Initializing PulseBlasterDDS ---\n");

// Check for proper initialization of PulseBlasterDDS

// Set clock frequency

set_clock(100);

// Start programming the frequency registers

start_programming(FREQ_REGS);

// Program the registers in order from 0 to 15

// Valid units are MHz, kHz, and Hz - default is MHz

set_freq(1*MHz); // Set register 0

set_freq(2*MHz); // Set register 1

set_freq(3*MHz); // Set register 2

set_freq(4*MHz); // Set register 3

set_freq(5*MHz); // Set register 4

set_freq(6*MHz); // Set register 5

set_freq(7*MHz); // Set register 6

set_freq(8*MHz); // Set register 7

set_freq(9*MHz); // Set register 8

set_freq(10*MHz); // Set register 9

set_freq(11*MHz); // Set register 10

set_freq(12*MHz); // Set register 11

set_freq(13*MHz); // Set register 12

set_freq(14*MHz); // Set register 13

set_freq(15*MHz); // Set register 14

set_freq(16*MHz); // Set register 15

// Start programming the phase registers for DAC_OUT_1

start_programming(PHASE_REGS_0);

// Program the registers in order from 0 to 15

// Units are in degrees

set_phase(0); // Set register 0

set_phase(22.5); // Set register 1

set_phase(45); // Set register 2

set_phase(67.5); // Set register 3

set_phase(90); // Set register 4

set_phase(112.5); // Set register 5

set_phase(135); // Set register 6

set_phase(157.5); // Set register 7

set_phase(180); // Set register 8

set_phase(202.5); // Set register 9

set_phase(225); // Set register 10

set_phase(247.5); // Set register 11

set_phase(270); // Set register 12

set_phase(292.5); // Set register 13

set_phase(315); // Set register 14

set_phase(337.5); // Set register 15

// Start programming the phase registers for DAC_OUT_0 and DAC_OUT_2

start_programming(PHASE_REGS_1);

// Program the registers in order from 0 to 15

// Units are in degrees

set_phase(0); // Set register 0

set_phase(22.5); // Set register 1

set_phase(45); // Set register 2

set_phase(67.5); // Set register 3

set_phase(90); // Set register 4

set_phase(112.5); // Set register 5

set_phase(135); // Set register 6

set_phase(157.5); // Set register 7

set_phase(180); // Set register 8

set_phase(202.5); // Set register 9

set_phase(225); // Set register 10

set_phase(247.5); // Set register 11

set_phase(270); // Set register 12

set_phase(292.5); // Set register 13

set_phase(315); // Set register 14

set_phase(337.5); // Set register 15

//Begin pulse program

start_programming(PULSE_PROGRAM);

int start, loop, sub; // define instruction labels

sub = 5; // Since we are going to jump forward in our program, we need to

// define this variable by hand.  Instructions start at 0 and count up

// Instruction 0 - Jump to Subroutine at Instruction 5 in 1us

start =	pb_inst(0,0,TX_ANALOG_ON,0,RX_ANALOG_ON,0x3FF,JSR,sub,1*us);

// Instruction 1 - Beginning of Loop (Loop 3 times).  Continue to next instruction in 1us

loop =	pb_inst(0,0,TX_ANALOG_OFF,0,RX_ANALOG_OFF,0x0,LOOP,3,1*us);

// Instruction 2 - End of Loop.  Return to beginning of loop or continue to next instruction in 1us

pb_inst(0,0,TX_ANALOG_ON,0,RX_ANALOG_OFF,0x0,END_LOOP,loop,1*us);

// Instruction 3 - Stay here for (5*1us) then continue to Instruction 4

pb_inst(0,0,TX_ANALOG_OFF,0,RX_ANALOG_ON,0x0,LONG_DELAY,5,1*us);

// Instruction 4 - Branch to "start" (Instruction 0) in 1us

pb_inst(0,0,TX_ANALOG_OFF,0,RX_ANALOG_ON,0x0,BRANCH,start,1*us);

// Instruction 5 - Continue to next instruction in 2us

pb_inst(0,4,TX_ANALOG_OFF,8,RX_ANALOG_ON,0x0,CONTINUE,0,2*us);

// Instruction 6 - Return from Subroutine to Instruction 1 in 2us

pb_inst(0,8,TX_ANALOG_OFF,8,RX_ANALOG_ON,0x0,RTS,0,2*us);

// End of programming registers and pulse program

stop_programming();

// Start execution of the pulse program

start_pb();

// Release control of the PB board

pb_close();

}

# Appendix II: Programming the PulseBlasterDDS Using Direct Outputs

If you do not wish to use the provided C functions, you can take advantage of the PBD03PC_outp() function in your own applications.  An explanation of using this function to program the board is included below.

## Using DLL Functions to Send Instructions

The provided driver DLL, pbd03pc.dll, provides three functions necessary for programming the PulseBlasterDDS:

int PBD03PC_Init();

This function locates and initializes the PulseBlasterDDS PCI board.  It returns a 0 upon successful completion, or a negative number for an error.

int PBD03PC_outp(unsigned short address, int data);

This function sends 8 bits of data to the PCI board, at a specified address offset.  (Only the lowest 8 bits of the integer value are sent)

int PBD03PC_Close();

This function releases control of the PulseBlasterDDS PCI board.  It returns a 0 upon successful completion, or a negative number for an error.

### Building Instructions Using the DLL Functions

- To send instructions to the PulseBlasterDDS, the programmer must first call the PBD03PC_Init() function to locate and initialize the PulseBlasterDDS for sending instructions

- Instructions are then build using the PBD03PC_outp(unsigned short address, int data) function, 8 bits at a time.  For example, 10 PBD03PC_outp function calls are made to construct one 80-bit instruction in accordance with the Instruction Set Architecture

- Finally, the PBD03PC_Close() instruction is called to signal an end to programming, and to release control of the PulseBlasterDDS

## Programming Information

Initialization of the PulseBlasterDDS for operation involves a minimum of four steps.  The steps are as follows:

- Send LOAD NUMBER OF BYTES PER WORD instruction.

- Send SELECT PERIPHERAL DEVICE instruction.

- Send CLEAR ADDRESS COUNTER instruction.

- Loading data to memory.*

- Send PROGRAMMING FINISHED instruction.

* Once the board has been programmed, this step is optional.  If you are just restarting the pulse program after a STOP command or an HARDWARE_RESET has stopped the program, you can skip this step during re-initialization.

If these five commands are not sent from a PC, the PulseBlasterDDS will not run as desired.  All five instructions are required as an attempt to ensure that the device has been programmed before it can be armed.  The first time the system is used, the loading of the memory with data has to be performed.  Upon reset, four instructions must be executed to restart the device again.

WE for Peripherals: This register is used to select the peripheral that is to be programmed.  The value of this register that is used to select program memory is always zero and this is the default value for the register.  A complete listing of the values and the associated hardware that can be programmed when appropriately set is listed below in Table A1.1.  By selecting 0xFF for this value, no device is selected and the initial value of the output flags may be set.  (*Note that the clock signal to the FF must still be transitioned, which may be accomplished by writing to the base output port + 5 twice in succession.)

Table A1.1: Peripheral List

CLEAR ADDRESS COUNTER: The Address Counter is used to manufacture the memory address.  The Address Counter is not loadable; it can only be cleared and started at zero.  It is not possible to load a particular section of memory.  All loads must start from either the beginning of memory, or wherever the Address Counter left off.

Flag Initialization Strobe: The output flags of the PulseBlasterDDS can be programmed while the device is in a reset state.  This is useful to initialize flags after powering-up and to reset flags to a known state if a program must be aborted.  Writing to the Flag Initialization Strobe register will toggle the line used to clock data into the output latches.

LOAD_MEMORY: This instruction is used to specify data that should be used to program the memory used by the device.  Since the incoming data is taken only one byte at a time, the IBC must reconstruct the data word to be programmed.  The data word is reconstructed in the IBC most significant byte first.

PROGRAMMING FINISHED: This instruction enables the pattern generator of the PulseBlasterDDS.  This instruction prevents the pattern generator from accepting a hardware trigger or software start command before the device has been programmed.  Once the design has been programmed, the PROGRAMMING FINISIHED command must be sent to arm the device for operation.  After the pattern generator has been armed, any hardware trigger or software start command will cause the system to start operation. The PulseBlasterDDS can be reset by issuing the DEVICE_RESET command.  This will internally clear the PROGRAMMING FINISHED instruction and prevent the pattern generator from operating again until the IBC has been re-initialized.

In order to select each of the commands mentioned about, you write to the port base + offset.  A table of each offset’s meaning is included below, see next page.

| Offset | Command | Function |
| --- | --- | --- |
| 0 | DEVICE_RESET | Stops Pulse Program |
| 1 | DEVICE_START | Starts Pulse Program (only when in initialized state) |
| 2 | SELECT_BPW | Selects number of BPW (10 for instructions) |
| 3 | SELECT_DEVICE | Using Table 2, selects internal device to be programmed |
| 4 | CLEAR_ADDRESS_COUNTER | Resets internal memory address counter |
| 5 | FLAG_STROBE | Strobes output clock signal to preset digital outputs |
| 6 | DATA_TRANSFER | Data to be written to internal device |
| 7 | PROGRAMMING_FINISHED | Sets device in initialized state |

Table A1.2: Port Offset Command Functions

## Example Program

The following is an example of the output sequence to program the PulseBlasterDDS.  You must use the PBD03PC_outp(addr, data) function from the provided DLL in order to use this method.

//	Initialization:

PBD03PC_outp(0,0);			//(Issue device reset)

PBD03PC_outp(2,4);			//(Select number of bytes per word)

PBD03PC_outp(3,0xFF);			//(Select device to program (Flag initial values))

PBD03PC_outp(4,0);			//(Reset address counter)

//	Set initial flag values

//	Values for this example are "0x000000f0"

PBD03PC_outp(6,0);			//(Data transfer)

PBD03PC_outp(6,0);			//(Data transfer)

PBD03PC_outp(6,0);			//(Data transfer)

PBD03PC_outp(6,0xF0);			//(Data transfer)

PBD03PC_outp(5,0);			//(Clock data into external buffer)

PBD03PC_outp(5,0);			//(Return clock signal to low)

//	Set up DDS frequency registers

PBD03PC_outp(0,0);			//(Issue device reset)

PBD03PC_outp(2,4);			//(Select number of bytes per word)

PBD03PC_outp(3,1);			//(Select device to program (DDS Frequency Registers))

PBD03PC_outp(4,0);			//(Reset address counter)

/*

DDS Register Values

Reg0 = 051EB852 (1 MHz)

Reg1 = 0A3D70A4 (2 MHz)

.

.

Reg15 = 147AE148 (4 MHz)

Formula for finding these values:

REG0 = DESIRED_FREQUENCY * 232 / PBDDS_CLOCK

= 1 MHz * 232 / 50 MHz = 858993459.2 = 0x051EB852

*/

PBD03PC_outp(6,0x05);		//(Data Transfer - Byte 3 of Reg0)

PBD03PC_outp(6,0x1E);		//(Data Transfer - Byte 2 of Reg0)

PBD03PC_outp(6,0xB8);		//(Data Transfer - Byte 1 of Reg0)

PBD03PC_outp(6,0x52);		//(Data Transfer - Byte 0 of Reg0)

PBD03PC_outp(6,0x05);		//(Data Transfer - Byte 3 of Reg1)

PBD03PC_outp(6,0x3D);		//(Data Transfer – Byte 2 of Reg1)

PBD03PC_outp(6,0x70);		//(Data Transfer - Byte 1 of Reg1)

PBD03PC_outp(6,0xA4);		//(Data Transfer - Byte 0 of Reg1)

.

.

PBD03PC_outp(6,0x14);		//(Data Transfer - Byte 3 of Reg15)

PBD03PC_outp(6,0x7A);		//(Data Transfer - Byte 2 of Reg15)

PBD03PC_outp(6,0xE1);		//(Data Transfer - Byte 1 of Reg15)

PBD03PC_outp(6,0x48);		//(Data Transfer - Byte 0 of Reg15)

//	Set up phase registers

PBD03PC_outp(0x0);			//(Issue device reset)

PBD03PC_outp(2,4);			//(Select number of bytes per word)

PBD03PC_outp(3,2);			//(Select device to program (TX Phase Registers))

PBD03PC_outp(4,0);			//(Reset address counter)

/*

Phase Register Values

Reg0 = 093E93E9 (13 degrees)

Reg1 = 40000000 (90 degrees)

.

.

Reg15 = E0000000(315 degrees)

Formula for finding these values:

REG0 = DESIRED_PHASE * 232 / 360

= 13 * 232 / 360 = 155096041.24  0x093E93E9

* Note that only the top 3 hex digits (0x093) are actually used in programming the phase register.  The other digits must still be sent so that the DDS word remains 32 bits long.

*/

PBD03PC_outp(6,0x09);		//(Data Transfer - Byte 3 of Reg0)

PBD03PC_outp(6,0x3E);		//(Data Transfer - Byte 2 of Reg0)

PBD03PC_outp(6,0x93);		//(Data Transfer - Byte 1 of Reg0)

PBD03PC_outp(6,0xE9);		//(Data Transfer - Byte 0 of Reg0)

PBD03PC_outp(6,0x40);		//(Data Transfer - Byte 3 of Reg1)

PBD03PC_outp(6,0x00);		//(Data Transfer - Byte 2 of Reg1)

PBD03PC_outp(6,0x00);		//(Data Transfer - Byte 1 of Reg1)

PBD03PC_outp(6,0x00);		//(Data Transfer - Byte 0 of Reg1)

.

.

PBD03PC_outp(6,0xE0);		//(Data Transfer - Byte 3 of Reg15)

PBD03PC_outp(6,0x00);		//(Data Transfer - Byte 2 of Reg15)

PBD03PC_outp(6,0x00);		//(Data Transfer - Byte 1 of Reg15)

PBD03PC_outp(6,0x00);		//(Data Transfer - Byte 0 of Reg15)

//	Pulse Program Setup

PBD03PC_outp(0,0);			//(Issue Device Reset)

PBD03PC_outp(2,10);		//(Select number of bytes per word)

PBD03PC_outp(3,0);			//(Select device to program (RAM))

PBD03PC_outp(4,0);			//(Reset address counter)

//Send pulse program

PBD03PC_outp(6,0x18);		//(Byte 9 of first instruction)

PBD03PC_outp(6,0xFF);		//(Byte 8 of first instruction)

PBD03PC_outp(6,0xFF);		//(Byte 7 of first instruction)

PBD03PC_outp(6,0x00);		//(Byte 6 of first instruction)

PBD03PC_outp(6,0x00);		//(Byte 5 of first instruction)

PBD03PC_outp(6,0x00);		//(Byte 4 of first instruction)

PBD03PC_outp(6,0x00);		//(Byte 3 of first instruction)

PBD03PC_outp(6,0x00);		//(Byte 2 of first instruction)

PBD03PC_outp(6,0x00);		//(Byte 1 of first instruction)

PBD03PC_outp(6,0x07);		//(Byte 0 of first instruction)

PBD03PC_outp(6,0xXX);		//(Byte 9 of second instruction)

/*

Continue this process for all instructions.  The

explanation of how to create the 80 bit instruction

words is included above.  When finished with all

instructions, continue with the sequence below.

*/

PBD03PC_outp(7,0);			//(Programming Finished)

/*

Only execute the following command when you are ready

for the program to start running.

*/

PBD03PC_outp(1,0);			//(Start pulse program)

# Contact Information

Phone				(352) 271-7383

FAX					(352) 371-8679

Email					sales@spincore.com

Web					http://www.spincore.com/

Product URL:

http://www.pulseblaster.com/CD/PulseBlasterDDS/PCI/SP3

Note that the usefulness of a waveform with two samples per period is limited, and, depending on applications, practical considerations would often call for more than two samples per period.

These functions and library files have been generated and tested with the MS Visual Studio 6 environment.  Support and updated functions/DLLs for other environments may be provided upon request if available.