PulseBlasterESR-Pro-cPCI™

Owner’s Manual

SpinCore Technologies, Inc.

http://www.spincore.com

Congratulations and thank you for choosing a design from SpinCore Technologies, Inc.

We appreciate your business!

At SpinCore we try to fully support the needs of our customers.  If you are in need of assistance, please contact us and we will strive to provide the necessary support.

© 2000-2005 SpinCore Technologies, Inc. All rights reserved.
SpinCore Technologies, Inc. reserves the right to make changes to the product(s) or information herein without notice. PulseBlasterESR™, PulseBlaster™, SpinCore, and the SpinCore Technologies, Inc. logos are trademarks of SpinCore Technologies, Inc. All other trademarks are the property of their respective owners.

SpinCore Technologies, Inc. makes every effort to verify the correct operation of the equipment.  This equipment version is not intended for use in a system in which the failure of a SpinCore device will threaten the safety of equipment or person(s).

Table of Contents

I. Introduction	4

Product Overview	4

Board Architecture	5

Block Diagram	5

Output signals	5

Timing characteristics	5

Instruction set	6

External triggering	6

Summary	6

Specifications	6

TTL Specifications	6

Pulse Parameters	6

Pulse Program Control Flow (Common)	6

II. Installation	7

Installing the PulseBlaster Driver for Windows 98/XP Operating System	7

For Windows 2000/XP	7

Testing Control of the PulseBlasterESR-Pro	8

III. Programming the PulseBlasterESR-Pro	9

Instruction Set Architecture	9

Machine-Word Definition	9

Breakdown of 80-bit Instruction Word	9

Using C Functions to Program the PulseBlaster	11

Example Use of C Functions	13

IV. Connecting to the PulseBlasterESR-Pro Board	14

Connector Information	14

BNC Header Connectors	14

IDC Connector Flag 0..11_Out Pin Assignments	14

IDC Connector Flag 12..23_Out Pin Assignments	15

IDC Connector Flag 24..35_Out Pin Assignments	15

Header HWTrig/Reset	16

V. Using PulseBlasterESR-Pro under Linux	17

General Information	17

Appendix II: Programming the PulseBlaster Using Direct Outputs	17

Using DLL Functions to Send Instructions	17

Building Instructions Using the DLL Functions	17

Programming Information	18

Example Program	20

Contact Information	22

# I. Introduction

## Product Overview

The PulseBlasterESR-PRO-cPCI device is a high-speed, multi-channel, intelligent pulse/pattern generation unit in the Compact-PCI (c-PCI) form factor.   The digital logic output signals are 3.3 V TTL compatible.

The intelligence of the PulseBlasterESR-Pro comes from an imbedded microprogrammed controller core (uPC).  The controller is able to execute instructions that allow it to control program flow.  This means that the PulseBlasterESR-Pro understands Operational Control Codes, OpCodes, and will execute them much the same way as a general-purpose microprocessor does.  The PulseBlasterESR-Pro’s microcontroller is different from the general-purpose microprocessor in that it does not contain an arithmetic logic unit (ALU) and is, therefore, incapable of doing mathematical or logical calculations.   However, a unique and distinguishing feature of the PulseBlasterESR-Pro processor is that the execution time of instructions is user programmable.  This feature makes the PulseBlasterESR-Pro processor capable of executing complex patterns at greatly varying update rates, ranging from nanoseconds to years, with a constant setting accuracy of just one clock period (e.g., a 2.5 ns setting accuracy when using a 400 MHz clock frequency).

## Board Architecture

### Block Diagram

Figure 1 presents the general architecture of the PulseBlasterESR-Pro system.  The major building blocks are the SRAM memory, the microcontroller (uPC), the integrated bus controller (IBC), the counter, and the output buffers.  The entire logic design, including the SRAM memory and output buffers, is contained on a single silicon chip, making it a System-on-a-Chip design. User control to the system is provided through the integrated bus controller (IBC) over the c-PCI bus.

PCI Bus

Figure 1: PulseBlasterESR-Pro board architecture

### Output signals

The PulseBlasterESR-Pro allows for 21 digital output signal lines.  All 21 signal lines are routed to two sets of 26-pin IDC on-board connectors.  The first four output bits are also routed to four bracket mounted BNC connectors.  The output signals are impedance matched to 50 ohm.

The 21 individually controlled digital output bits comply with the TTL-levels’ standard, and are capable of delivering 25 mA per bit/channel.  If more output current is necessary, the individual bits/channels can be driven in parallel.

### Timing characteristics

PulseBlaster’s timing controller accepts an internal (on-board) crystal oscillator up to 150 MHz.  The input frequency is internally doubled, or tripled if necessary, up to 400 MHz.  The innovative architecture of the timing controller allows the processing of either simple timing instructions (delays of up to 28 = 256 clock cycles), or double-length timing instructions (up to 228 clock cycles long – 0.89 seconds at 300 MHz).  Regardless of the type of timing instruction, the timing resolution remains constant for any delay – just one clock period (e.g., 3.3 ns at 300 MHz).

The core-timing controller has a very short minimum delay cycle – only six clock periods.  This translates to a 15 ns pulse/delay/update at 400 MHz.  The PulseBlasterESR-Pro is also capable of generating pulses on the four BNC connectors of lengths one clock period, two clock periods, three clock periods, four clock periods, four and one half clock periods, five clock periods, and five and one half clock periods.

### Instruction set

PulseBlaster’s design features a set of commands for highly flexible program flow control. The micro-programmed controller allows for programs to include branches, subroutines, and loops at up to 8 nested levels – all this to assist the user in creating dense pulse programs that cycle through repetitious events, especially useful in numerous multidimensional spectroscopy and imaging applications.

### External triggering

PulseBlasterESR-Pro can be triggered and/or reset externally via dedicated hardware lines.  The two separate lines combine the convenience of triggering (e.g., in cardiac gating) with the safety of the "stop/reset" line.  The required control signals are “active low” (or short to ground).

### Summary

PulseBlasterESR-Pro is a versatile, high-performance pulse/pattern TTL signal generator operating at speeds of up to 400 MHz and capable of generating pulses ranging from 2.5ns to .89s per instruction and delays/intervals ranging from 15 ns to 0.89 s per instruction. It can accommodate pulse programs with highly flexible control commands of up to 8192 program words. Its high-current output logic bits are independently controlled with an output voltage of 3.3 V  (unterminated).

## Specifications

### TTL Specifications

- 21 individually controlled digital output lines (TTL levels, 3.3 V logical “one”)

- 4 bracket mounted BNC connectors, impedance matched to 50 ohm.

- variable pulses/delays for every TTL line

- 25 mA output current per TTL line

### Pulse Parameters

- 2.5 ns shortest pulse (at 400 MHz)

- 15 ns shortest interval (at 400 MHz)

- 0.89 s longest pulse/interval (at 400 MHz)

- 2.5 ns pulse/interval resolution (at 400 MHz)

- up to 8192 instructions

- external  triggering and reset – TTL levels

### Pulse Program Control Flow (Common)

- loops, nested 8 levels deep

- 20 bit loop counters (max. 1,048,576 repetitions)

- subroutines, nested 8 levels deep

- wait for trigger – 8 clock cycle latency (26.6ns at 300 MHz), adjustable to 0.89 seconds in duration

- 15 MHz max. re-triggering frequency

# II. Installation

## Installing the PulseBlaster Driver for Windows 98/XP Operating System

- Go to http://www.pulseblaster.com/CD/spinapi/ and download spinapi.zip.

- Unzip the files to the their own directory and make a note of this directory.  It is convenient to use the desktop.  For Windows 98, this directory will be c:\windows\desktop\spinapi.  For Windows XP/2000, this will probably be c:\documents and settings\username\desktop\spinapi, where username is your username used to log in to Windows.

- Turn off your computer.

- Insert the PulseBlaster board into an empty PCI slot.

- Turn on your computer.

For Windows 98

- After booting, an “Add New Hardware Wizard” dialog box will appear.  Click the Next Button.

- Select Search for the best driver for your device (Recommended) and click the Next button.

- Select Specify a location and click the Browse… button.

- Browse to the folder you created when downloading spinapi.zip and browse to the /dll/ directory.

- You will return to the previous screen.  Now, click the Next button.

- Windows is now ready to install the PulseBlaster driver.  Click the Next button.

- Windows will now copy the necessary files to your PC.  When the process completes, click Finish.

- Browse to the folder you created when downloading spinapi.zip and browse to the /dll/ directory.

- Run install_wd.bat by double clicking on install_wd.bat.

## For Windows 2000/XP

Steps 1 to 5: as described earlier.

6.  After booting, the “Found New Hardware Wizard” should appear.  Choose “Install from a list or specific location” and click Next

7.  Choose “Include this location in the search” and browse to the directory you unzipped the spinapi.zip to.  Browse to the /dll/ directory.  Click next.

- While windows installs the driver, a “Files Needed” dialog may pop up.  Choose the directory you unzipped the drivers to, browse to the /dll/ directory, and click ok.

- Browse to the folder you created when downloading spinapi.zip and browse to the /dll/ directory.

- Run install_wd.bat by double clicking on install_wd.bat.

## Testing Control of the PulseBlasterESR-Pro

- Open an MS-DOS window.  You can do this by clicking start … run.  Then, for windows 98 type command and press enter.  For windows XP and 2000 type cmd and press enter.

- Navigate to the directory you unzipped spinapi.zip to.  You can do this by typing cd\dir where dir=the directory you unzipped spinapi.zip to in step 2 of the installation above.

- Navigate to spinapi/PBESR-PRO by typing ‘cd PBESR-PRO’.

- Type looptest and press enter.  looptest.exe assumes you have a 300 MHz clock.  You can change this by typing looptest x, where x= your clock speed.  Looptest will output pulses on all 4 BNC connectors.  Observe the output of looptest.exe in the MS-DOS window for details.

- Type subtest and press enter.  subtest.exe assumes you have a 300 MHz clock.  You can change this by typing subtest x, where x= your clock speed. subtest will output pulses on all 4 BNC connectors.  Observe the output of subtest.exe in the MS-DOS window for details.

- Type longdelaytest and press enter.  longdelaytest.exe assumes you have a 300 MHz clock.  You can change this by typing longdelaytest x, where x= your clock speed. longdelaytest will output pulses on all 4 BNC connectors.  Observe the output of longdelaytest.exe in the MS-DOS window for details.

- Type short_pulse_ex and press enter.  Short_pulse_ex will output a train of pulses, with the pulse high for one clock cycle, low for 9 clock cycles.

The PulseBlaster board is now ready for use

# III. Programming the PulseBlasterESR-Pro

## Instruction Set Architecture

### Machine-Word Definition

The PulseBlaster pulse timing and control processor implements an 80-bit wide Very Long Instruction Word (VLIW) architecture. The VLIW memory words have specific bits/fields dedicated to specific purposes, and every word should be viewed as a single instruction of the micro-controller. The maximum number of instructions that can be loaded to on-board memory is 4096. The execution time of instructions can be varied and is under (self) control by one of the fields of the instruction word – the shortest being five clock cycles and the longest being 2^28 clock cycles.  All instructions have the same format and bit length, and all bit fields have to be filled.  Figure 2 shows the fields and bit definitions of the 80-bit instruction word.

Bit Definitions for the 80-bit Instruction Word (VLIW)

Output/Control Word   |    Data Field    |     OP Code    |   Delay Count

(24 bits)             (20 bits)          (4 bits)         (32 bits)

Figure 2: Bit definitions of the 80-bit instruction/memory word

### Breakdown of 80-bit Instruction Word

The 80-bit VLIW is broken up into 4 sections

1. Output Pattern and Control Word - 24 bits

2. Data Field - 20 bits

3. OP Code - 4 bits

4. Delay Count - 32 bits; all bits above the eighth bit must be programmed with zeros.

#### Output Pattern and Control Word

Please refer to Table 1 for output pattern and control bit assignments of the 21-bit output/control word.

| Bit # |  | Bit # | Function |
| --- | --- | --- | --- |
| 23 | Controls Pulse Length for BNC connectors | 11 | Output Connector labeled Flag0..11 Out pin 23 |
| 22 | Controls Pulse Length for BNC connectors | 10 | Output Connector labeled Flag0..11 Out pin 21 |
| 21 | Controls Pulse Length for BNC connectors | 9 | Output Connector labeled Flag0..11 Out pin 19 |
| 20 | Output Connector labeled Flag0..11 Out pin 17 | 8 | Output Connector labeled Flag0..11 Out pin 17 |
| 19 | Output Connector labeled Flag0..11 Out pin 15 | 7 | Output Connector labeled Flag0..11 Out pin 15 |
| 18 | Output Connector labeled Flag0..11 Out pin 13 | 6 | Output Connector labeled Flag0..11 Out pin 13 |
| 17 | Output Connector labeled Flag0..11 Out pin 11 | 5 | Output Connector labeled Flag0..11 Out pin 11 |
| 16 | Output Connector labeled Flag0..11 Out pin 9 | 4 | Output Connector labeled Flag0..11 Out pin 9 |
| 15 | Output Connector labeled Flag0..11 Out pin 7 | 3 | Output Connector labeled Flag0..11 Out pin 7 |
| 14 | Output Connector labeled Flag0..11 Out pin 5 | 2 | Output Connector labeled Flag0..11 Out pin 5 |
| 13 | Output Connector labeled Flag0..11 Out pin 3 | 1 | Output Connector labeled Flag0..11 Out pin 3 |
| 12 | Output Connector labeled Flag12..23 Out pin 1 | 0 | Output Connector labeled Flag0..11 Out pin 1 |

Table 1: Output Pattern and Control Word Bits

When the bit corresponding to an IDC output connector is one, it will remain high for the duration of the instruction.  When the bit is zero, it will remain low for the duration of the instruction.

#### BNC Short Pulses

Please refer to Table 2 for pulse lengths on the BNC connectors.  The first column represents values on bits 21-23 of the Output/Control word and the second column represents the delay count.  The third column represents the number of clock periods that the instruction is high.  The fourth column assumes a clock of 300MHz resulting in a 3.33 ns period.

| Bits 21-23 | Delay Count | Clock periods | Pulse Length (ns) |
| --- | --- | --- | --- |
| 1 | 2 | 1 | 3.33 |
| 2 | 2 | 2 | 6.66 |
| 3 | 2 | 3 | 9.99 |
| 4 | 2 | 4 | 13.32 |
| 5 | 2 | 4.5 | 14.98 |
| 5 | 3 | 5 | 16.65 |
| 6 | 3 | 5.5 | 18.31 |
| 6 | 4 | 6.5 | 21.64 |
| 6 | 5 | 7.5 | 24.97 |
| ... | ... | ... | ... |

Table 2: BNC short Pulses

When bits 21-23 are zero, the BNC connectors remain low for the duration of the instruction.  When they are between 1 and 6, the BNCs pulse goes high at the beginning of the instruction.  When bits 21-23 are 6, the pulse length is (delay value + 2.5) * (clock period).  The instruction length is (delay value + 4) * (clock period) so there will always be a 1.5 clock period interval after the pulse goes low and before the next instruction begins on the BNCs.

#### Data Field and Op Code

Please refer to Table 2 for information on the available operational codes (OpCode) and the associated data field functions (the data field's function is dependent on the Op Code)

| Op Code # | Inst | Inst_data | Function |
| --- | --- | --- | --- |
| 0 | CONTINUE | Ignored | Program execution continues to next instruction |
| 1 | STOP | Ignored | Stop execution of program (*Note all TTL values remain from previous instruction, and analog outputs turn off) |
| 2 | LOOP | Number of desired loops. This value must be greater than or equal to 1. | Specify beginning of a loop. Execution continues to next instruction. Data used to specify number of loops |
| 3 | END_LOOP | Address of beginning of loop | Specify end of a loop. Execution returns to begging of loop and decrements loop counter. |
| 4 | JSR | Address of first subroutine instruction | Program execution jumps to beginning of a subroutine |
| 5 | RTS | Ignored | Program execution returns to instruction after JSR was called |
| 6 | BRANCH | Address of next instruction | Program execution continues at specified instruction |
| 7 | LONG_DELAY | Number of desired loops. This value must be greater than or equal to 2. | For long interval instructions. Data field specifies a multiplier of the delay field. Execution continues to next instruction |
| 8 | WAIT | Ignored | Program execution stops and waits for software or hardware trigger. Execution continues to next instruction after receipt of trigger. The latency is equal to the delay value entered in the WAIT instruction line plus a fixed delay of 6 clock cycles. |

Table 3: Op Code and Data Field Description

#### Delay Count

The value of the Delay Count field (a 32-bit value) determines how long the current instruction should be executed.  The allowed minimum value of this field is 0x00000002 and the allowed maximum is 0x000000FF.  The timing controller has a fixed delay of four clock cycles and the value that one enters into the Delay Count field should account for this inherent delay.

## Using C Functions to Program the PulseBlaster

A series of functions have been written to control the board and facilitate the construction of pulse program instructions.  It should be noted that the pb_inst C function accepts any delay value greater than 17 ns.  Since the clock period is 3.3 ns for a 300 MHz clock, values which are not integer multiples of the clock will be rounded to the closest integer multiple.  For more precise control for short pulses, examine the example program in Appendix II Programming the PulseBlaster Using Direct Outputs.

In order to use these functions, the DLL (spinapi.dll), the library file (libspinapi.a for mingw, spinapilibgcc for borland, and spinapi.lib for msvc), the header file (spinapi.h), must be in the working directory of your C compiler.

int pb_init();

Initializes PulseBlaster board.  Needs to be called before calling any functions using the PulseBlaster.  Returns a negative number on an error or 0 on success.

int pb_close();

Releases PulseBlaster board.  Needs to be called as last command in pulse program.  Returns a negative number on an error or 0 on success.

void set_clock(double clock_freq);

Used to set the clock frequency of the board.  The variable clock_frequency is specified in MHz when no units are entered.  Valid units are MHz, kHz, and Hz.  The default clock value is 50MHz.  You only need to call this function if you are not using a –50 board.

int start_programming(int device);

Used to initialize the system to receive programming information.  It accepts a parameter referencing the target for the instructions.  The only alid value for device is PULSE_PROGRAM, It returns a 0 on success or a negative number on an error.

int pb_inst(int flags, int inst, int inst_data, double length);

Used to send one instruction of the pulse program.  Should only be called after start_programming(PULSE_PROGRAM) has been called.  It returns a negative number on an error, or the instruction number upon success.  If the function returns –99, an invalid parameter was passed to the function.  Instructions are numbered starting at 0.

int flags – determines state of each TTL output bit.  Valid values are 0x000000 to 0x1FFFFF.  For 	example, 0x000010 would correspond to bit 5 being on, and all other bits being off.

int inst – determines which h type of instruction is to be executed. Please see Table 3 for details.

int inst_data – data to be used with the previous inst field.  Please see Table 3 for details.

int length – duration of this pulse program instruction, specified in ns.

int stop_programming();

Used to tell that programming the board is complete.  Board execution cannot start until this command is received.  It returns a 0 on success or a negative number on an error.

int start_pb();

Once board has been programmed, this instruction will start execution of pulse program.  It returns a 0 on success or a negative number on an error.

int stop_pb();

Stops output of board.  Analog output will return to ground, and TTL outputs will remain in the state they were in when stop command was received.  It returns a 0 on success or a negative number on an error.

### Example Use of C Functions

#include <stdio.h>

#define PBESRPRO

#include "spinapi.h"

#define CLOCK 300

int main(void)

{

int start;

set_clock(CLOCK);  // Set Because the Board Operates at 300MHz

if (pb_init() != 0)

{

printf("--- Error Initializing PulseBlaster ---\n");

return -1;

}

start_programming(PULSE_PROGRAM);	// Prepare the Board to Receive

// Pulse Program Instructions

// Instruction 0 - Continue to instruction 1 in 20ns

// Flags = 0x1FFFFF, OPCODE = CONTINUE

// ONE_PERIOD equals 1 for bits 23-21 ... other valid constants are

// TWO_PERIODS, THREE_PERIODS, FOUR_PERIODS, FIVE_PERIODS, and ON

start = pb_inst(ONE_PERIOD|0xF, CONTINUE, 0, 80*ns);

// Instruction 1 - Continue to instruction 2 in 40ns

// Flags = 0x1, OPCODE = CONTINUE

// inst length = 40 ns

// 40*ns length = delay value of 8 in instruction word

// pulse in BNC 0 high for (8 + 4) * 3.33 - 1.5 * 3.33 = 34.96

pb_inst(ON|0xF, CONTINUE, 0, 40*ns);

// Instruction 2 - Branch to "start" (Instruction 0) in 40ns

// 0x0, OPCODE = BRANCH, Target = start

pb_inst(0, BRANCH, start, 40*ns);

stop_programming();// Finished Sending Instructions

start_pb();//  Run the Program

pb_close();  // Release Control of the PulseBlasterESR board

return 0;

}

# IV. Connecting to the PulseBlasterESR-Pro Board

## Connector Information

### BNC Header Connectors

The four bracket mounted BNC connectors output the first four bits, starting from the list-significant bit, as depicted below.  When using coaxial transmission lines, please terminate the lines with a 50 ohm impedance.

Bit 3

Bit 2

Bit 1

Bit 0

### IDC Connector Flag 0..11_Out Pin Assignments

The IDC connector labeled Flag0..11_Out outputs TTL signals generated by the user’s Program.  Please consult the table below for pin assignments.

| Pin Assignments |  |  |  |
| --- | --- | --- | --- |
| Pin# |  | Pin# |  |
| 1 | Bit 0 | 14 | GND |
| 2 | GND | 15 | Bit 7 |
| 3 | Bit 1 | 16 | GND |
| 4 | GND | 17 | Bit 8 |
| 5 | Bit 2 | 18 | GND |
| 6 | GND | 19 | Bit 9 |
| 7 | Bit 3 | 20 | GND |
| 8 | GND | 21 | Bit 10 |
| 9 | Bit 4 | 22 | GND |
| 10 | GND | 23 | Bit 11 |
| 11 | Bit 5 | 24 | GND |
| 12 | GND | 25 | NC |
| 13 | Bit 6 | 26 | GND |

Table 4: Output bits and signals of the PulseBlaster

### IDC Connector Flag 12..23_Out Pin Assignments

The IDC connector labeled Flag12..23_Out outputs TTL signals generated by the user’s Program.  Please consult the table below for pin assignments.

| Pin Assignments |  |  |  |
| --- | --- | --- | --- |
| Pin# |  | Pin# |  |
| 1 | Bit 12 | 14 | GND |
| 2 | GND | 15 | Bit 19 |
| 3 | Bit 13 | 16 | GND |
| 4 | GND | 17 | Bit 20 |
| 5 | Bit 14 | 18 | GND |
| 6 | GND | 19 | NC |
| 7 | Bit 15 | 20 | GND |
| 8 | GND | 21 | NC |
| 9 | Bit 16 | 22 | GND |
| 10 | GND | 23 | NC |
| 11 | Bit 17 | 24 | GND |
| 12 | GND | 25 | NC |
| 13 | Bit 18 | 26 | GND |

Table 5: Output bits and signals of the PulseBlaster

### IDC Connector Flag 24..35_Out Pin Assignments

The IDC connector labeled Flag24..35_Out outputs TTL signals based on the status of the user’s Program.  Please consult the table below for pin assignments.

| Pin Assignments |  |  |  |
| --- | --- | --- | --- |
| Pin# |  | Pin# |  |
| 1 | Stopped | 14 | GND |
| 2 | GND | 15 | NC |
| 3 | Reset | 16 | GND |
| 4 | GND | 17 | NC |
| 5 | Running | 18 | GND |
| 6 | GND | 19 | NC |
| 7 | Waiting | 20 | GND |
| 8 | GND | 21 | NC |
| 9 | NC | 22 | GND |
| 10 | GND | 23 | NC |
| 11 | NC | 24 | GND |
| 12 | GND | 25 | NC |
| 13 | NC | 26 | GND |

Table 6: Status signals of the PulseBlaster

The status pins correspond to the current state of the pulse program and are defined as follows:

Stopped – Driven high when the PulseBlaster device has encountered a STOP Op Code during program execution and has entered a stopped state.

Reset – Driven low when the PulseBlaster device is in a RESET state and must be reprogrammed before code execution can begin again.

Running – Driven high when the PulseBlaster device is executing a program.  It is low when the PulseBlaster enters either a reset or idle state.

Waiting – the PulseBlaster device has encountered a WAIT Op Code and is waiting for the next trigger (either hardware or software) to resume operation.

### Header HWTrig/Reset

This is an input connector, for hardware triggering (HW_Trigger) and resetting (HW_Reset).

HW_Trigger is pulled high by default, and pin 2 is active (pin 4 = GND).  When a low value is detected (e.g., when shorting pins 2-4), it initiates code execution.   This trigger will also restart execution of a program from the beginning of the code if it is detected after the design has reached an idle state.  The idle state could have been created either by reaching the STOP Op Code of a program, or by the detection of the HW_Reset signal.  When the WAIT Op Code is used in the pulse program, the HW_Trigger will cause the program to continue to the next instruction.

HW_Reset is pulled high by default, and pin 1 is active (pin 3 = GND).  It can be used to halt the execution of a program by pulling it low (e.g., by shorting pins 1-3).  When the signal is pulled low during the execution of a program, the controller resets itself back to the beginning of the program.  Program execution can be resumed by either a software start command or by a hardware trigger.

# V. Using PulseBlasterESR-Pro under Linux

## General Information

The PulseBlasterESR-Pro boards can be used under the Linux operating systems.  Please contact SpinCore at sales@spincore.com for more information and latest drivers.  If necessary, SpinCore will provide, free of charge, the source code for the drivers.

# Appendix II: Programming the PulseBlaster Using Direct Outputs

If you do not wish to use the provided C functions, you can take advantage of the pb_outp() function in your own applications.  An explanation of using this function to program the board is included below.

## Using DLL Functions to Send Instructions

The provided driver DLL, spinapi.dll, provides three functions necessary for programming the PulseBlaster:

int pb_init();

This function locates and initializes the PulseBlaster PCI board.  It returns a 0 upon successful completion, or a negative number for an error.

int pb_outp(unsigned short address, unsigned char data);

This function sends 8 bits of data to the PCI board, at a specified address offset.

int pb_close();

This function releases control of the PulseBlaster PCI board.  It returns a 0 upon successful completion, or a negative number for an error.

### Building Instructions Using the DLL Functions

- To send instructions to the PulseBlaster, the programmer must first call the pb_init() function to locate and initialize the PulseBlaster for sending instructions

- Instructions are then build using the PB_outp(unsigned short address, int data) function, 8 bits at a time.  For example, 10 PB_outp function calls are made to construct one 80-bit instruction in accordance with the Instruction Set Architecture

- Finally, the PB_Close() instruction is called to signal an end to programming, and to release control of the PulseBlaster.

## Programming Information

Initialization of the PulseBlaster for operation involves a minimum of four steps.  The steps are as follows:

- Send LOAD NUMBER OF BYTES PER WORD instruction.

- Send SELECT PERIPHERAL DEVICE instruction.

- Send CLEAR ADDRESS COUNTER instruction.

- Loading data to memory.*

- Send PROGRAMMING FINISHED instruction.

* Once the board has been programmed, this step is optional.  If you are just restarting the pulse program after a STOP command or an HARDWARE_RESET has stopped the program, you can skip this step during re-initialization.

If these five commands are not sent from a PC, the PulseBlaster will not run as desired.  All five instructions are required as an attempt to ensure that the device has been programmed before it can be armed.  The first time the system is used, the loading of the memory with data has to be performed.  Upon reset, four instructions must be executed to restart the device again.

WE for Peripherals: This register is used to select the peripheral that is to be programmed.  The value of this register that is used to select program memory is always zero and this is the default value for the register.  A complete listing of the values and the associated hardware that can be programmed when appropriately set is listed below in Table A1.1.  By selecting 0xFF for this value, no device is selected and the initial value of the output flags may be set.  (*Note that the clock signal to the FF must still be transitioned, which may be accomplished by writing to the base output port + 5 twice in succession.)

Table A1.1: Peripheral List

CLEAR ADDRESS COUNTER: The Address Counter is used to manufacture the memory address.  The Address Counter is not loadable; it can only be cleared and started at zero.  It is not possible to load a particular section of memory.  All loads must start from either the beginning of memory, or wherever the Address Counter left off.

Flag Initialization Strobe: The output flags of the PulseBlaster can be programmed while the device is in a reset state.  This is useful to initialize flags after powering-up and to reset flags to a known state if a program must be aborted.  Writing to the Flag Initialization Strobe register will toggle the line used to clock data into the output latches.

LOAD_MEMORY: This instruction is used to specify data that should be used to program the memory used by the device.  Since the incoming data is taken only one byte at a time, the IBC must reconstruct the data word to be programmed.  The data word is reconstructed in the IBC most significant byte first.

PROGRAMMING FINISHED: This instruction enables the pattern generator of the PulseBlaster.  This instruction prevents the pattern generator from accepting a hardware trigger or software start command before the device has been programmed.  Once the design has been programmed, the PROGRAMMING FINISIHED command must be sent to arm the device for operation.  After the pattern generator has been armed, any hardware trigger or software start command will cause the system to start operation. The PulseBlaster can be reset by issuing the DEVICE_RESET command.  This will internally clear the PROGRAMMING FINISHED instruction and prevent the pattern generator from operating again until the IBC has been re-initialized.

In order to select each of the commands mentioned about, you write to the port base + offset.  A table of each offset’s meaning is included below.

| Offset | Command | Function |
| --- | --- | --- |
| 0 | DEVICE_RESET | Stops Pulse Program |
| 1 | DEVICE_START | Starts Pulse Program (only when in initialized state) |
| 2 | SELECT_BPW | Selects number of BPW (10 for instructions) |
| 3 | SELECT_DEVICE | Using Table 2, selects internal device to be programmed |
| 4 | CLEAR_ADDRESS_COUNTER | Resets internal memory address counter |
| 5 | FLAG_STROBE | Strobes output clock signal to preset digital outputs |
| 6 | DATA_TRANSFER | Data to be written to internal device |
| 7 | PROGRAMMING_FINISHED | Sets device in initialized state |

Table A1.2: Port Offset Command Functions

## Example Program

The following is an example of the output sequence to program the PulseBlaster.  You must use the pb_outp(addr, data) function from the provided DLL in order to use this method.

#include <stdio.h>

#define PBESRPRO

#include "spinapi.h"

int main()

{

//request and initialize control of PB

pb_init();

//tell the PB to start recieving pulse program instructions

//issue device reset

pb_outp(0,0);

//select number of bytes per word

pb_outp(2,10);

//select device to program (RAM)

pb_outp(3,0);

//reset address counter

pb_outp(4,0);

/////////////////////

//Bit pattern details

//bits 0 through 23 control the output of the four BNC connectors and

//the 21 IDC Flag outputs

//bits 0 through 3 control the output of the four BNC connectors and

//IDC flags 0 through 3

//bits 4 through 20 control the output of the IDC flags 4 through 20

//bits 21 through 23 select the length of the pulse for the BNC connectors

//for delay = 2

//1 on bits 23-21 results in a pulse high for 1 clock period

//2 on bits 23-21 results in a pulse high for 2 clock periods

//3 on bits 23-21 results in a pulse high for 3 clock periods

//4 on bits 23-21 results in a pulse high for 4 clock periods

//5 on bits 23-21 results in a pulse high for 4.5 clock periods

//for delay = 3

//5 on bits 23-21 results in a pulse high for 5 clock periods

//for delay >= 3

//6 on bits 23 through 21 results in a pulse high for:

//(delay+2.5)*(clock period)

//instruction length = (delay + 4) * clock period

////////////////////////////////

//Instruction 1

pb_outp(6,0x20);		//1 on bits 21-23 for pulse of length one clock cycle on BNCs

pb_outp(6,0x00);

pb_outp(6,0x0f);		//select BNC connectors 0-3

//opcode

pb_outp(6,0x00);

pb_outp(6,0x00);

pb_outp(6,0x00);		//opcode = CONTINUE

//delay

pb_outp(6,0x00);

pb_outp(6,0x00);

pb_outp(6,0x00);

pb_outp(6,0x02);		//delay = 2, inst length = (2+4)*(clock period)

////////////////////////////////

//Instruction 2

//pattern

pb_outp(6,0x00);

pb_outp(6,0x00);

pb_outp(6,0x00);

//opcode

pb_outp(6,0x00);		//data = 0

pb_outp(6,0x00);

pb_outp(6,0x06);		//opcode = BRANCH

//delay

pb_outp(6,0x00);

pb_outp(6,0x00);

pb_outp(6,0x00);

pb_outp(6,0x02);		//delay = 2, inst length = (2+4)*(clock period)

//end of programming

pb_outp(7,0);

//put device in run mode

pb_outp(1,0);

//relase control of PB

pb_close();

return 0;

}

# Contact Information

Email					sales@spincore.com

Web					http://www.spincore.com/

Product URL:

http://www.pulseblaster.com/CD/PulseBlasterESR/SP5

These functions and library files have been generated and tested with MinGW (www.mingw.com), Borland 5.5 (www.borland.com), MS Visual Studio 2003 (msdn.microsoft.com) compilers.